## Sheet
PERFORMANCE TEST
DUT
Manufacturer |  | Datron Instruments |  |  | S/N |  | 31255
Model Number |  | 4950 |  |  | Test date |  | 24-05-30 16:33:41
Meter Info |  | Datron 4950 STD |  |  | Ambient Temp. |  | 23.0
Options |  | 1,1,1 |  |  | CERT_AMB? |  | 23
Firmware |  | 890197-02.07 |  |  | TEMP? |  | 43.14
Date of Baseline Cal |  | "1999-SEP-16" |  |  | TEMP? BASE |  | 3.14
Date of Certified Cal |  | "2024-MAY-30" |  |  | TEMP? CERTIFIED |  | 0
CNFTST? |  | 0 |  |  | STANDBY? |  | "00:00"
DUT Confidence level |  | 30 Days Transfer Satbility
Calibrator
Reference standard |  | Options |  | Unc | Calibration Date |  | Due Date
Fluke 5700A |  |  |  | -- | -- |  | --
ADRmu12 |  | 10.000000 VDC |  | -- | -- |  | --
Fluke 742A-1 8927002 |  | 1.0000195 ohm |  | -- | -- |  | --
Fluke 742A-10k 7815006 |  | 10.000133 kohm
MFC
MFC Calibrate Date |  | 290524 |  |  | MFC last calibrated |  | 1 days
Calibrate Date WBFLAT |  | 100589 |  |  | MFC since WBFLAT |  | 0 days
Calibrate Date WB Gain |  | 250190 |  |  | MFC since WBGAIN |  | 0 days
MFC Calibrate Date Zero |  | 290524 |  |  | MFC since DCV Zero |  | 1 days
MFC Confidence level |  | 5700A 99% 90 Days |  |  | CAL CONST 11V Range +zero |  | 398.1793
CAL CONST 6.5V |  | +6.529890647135340E+00 |  |  | CAL CONST 11V Range -zero |  | 398.1784
CAL CONST 13V reference voltage |  | +1.306130764516422E+01 |  |  | CAL CONST, Zero calibration temperature |  | 23
CAL CONST 10KOHM true output resistance |  | +1.000045271113747E+04 |  |  | CAL CONST, ALL calibration temp |  | 23
CAL SONST 10kOHM standard resistance |  | +1.000069664541200E+04 |  |  | CAL_CONST? CHECK, WB_FLAT_TEMP |  | 23.9
CAL CONST 1 OHM |  | +9.996126071898848E-01 |  |  | CAL CONST,WBAND_TEMP
ETIME? |  | 6056 days 6.37 hours |  |  | CAL TEMP? CHECK
ONTIME? |  | 1 days 20.92 hours |  |  | CAL CONST,DACLIN |  | 0.359
DCV PERFORMANCE TEST
DCV Test VDC | Expected Value | Measured μV | Source Unc. μV | Lower Limit μV | Upper Limit μV | Deviation | DUT Transfer STB ppm | Test Result % of SPEC | MEAS SDEV μV
0.0mVDC | 0 | -0.01 | 7.5e-07 |  |  | N/A |  |  | 0.0097183 | 0.00631
0.0VDC | 0 | 4.2 | 7.5e-07 |  |  | N/A |  |  | 0 | 0.0117
0.0VDC | 0 | 15 | 7.5e-07 |  |  | N/A |  |  | 0 | 0.0797
0.0VDC | 0 | 230 | 7.5e-07 |  |  | N/A |  |  | 4.4096 | 0.917
0.0VDC | 0 | 0 | 7.5e-07 |  |  | N/A |  |  | 0 | 4.94
DCV Test VDC | DMM Range | Measured V | Source Unc. ppm | Lower Limit V | Upper Limit V | Measured ppm | DUT Transfer STB ppm | Test Result % of SPEC | MEAS SDEV ppm
0.01 | 100mV | 0.01 | 81.5 | [5 empty] | 0.78174 | 5.8e-07
0.02 | 100mV | 0.02 | 44 | [5 empty] | 0.48591 | 3.5e-07
0.05 | 100mV | 0.05 | 21.5 | [5 empty] | 0.19437 | 1.1e-07
0.1 | 100mV | 0.1 | 14 |  |  |  | 3 |  | 0.11304 | 5e-08
-0.01 | 100mV | -0.01 | -81.5 | [5 empty] | -1.118 | 5.7e-07
-0.02 | 100mV | -0.02 | -44 | [5 empty] | -0.43301 | 3.1e-07
-0.05 | 100mV | -0.05 | 21.5 | [5 empty] | -0.15635 | 1.1e-07
-0.1 | 100mV | -0.1 | 14 |  |  |  | 3 |  | -0.070711 | 5e-08
0.1 | 1V | 0.1 | 14 | [5 empty] | 0.99996 | 1.4e-07
0.2 | 1V | 0.2 | 10.25 | [5 empty] | 0.35355 | 8e-08
0.5 | 1V | 0.50001 | 5.9 | [5 empty] | 0.088191 | 4e-08
1 | 1V | 1 | 4.7 |  |  |  | 1.5 |  | 0.10929 | 2e-08
-0.1 | 1V | -0.099996 | 14 | [5 empty] | -0.70714 | 1.4e-07
-0.2 | 1V | -0.2 | 10.25 | [5 empty] | -0.16667 | 6e-08
-0.5 | 1V | -0.5 | 5.9 | [5 empty] | -0.088192 | 3e-08
-1 | 1V | -1 | 4.7 |  |  |  | 1.5 |  | -0.033333 | 2e-08
1 | 10V | 1 | 4.7 | [5 empty] | 0.7071 | 8e-08
2 | 10V | 2 | 4.1 | [5 empty] | 0.16667 | 4e-08
5 | 10V | 5 | 4.1 | [5 empty] | 0.088192 | 3e-08
10 | 10V | 10 | 3.8 |  |  |  | 1.5 |  | 0.044096 | 2e-08
19 | 10V | 19 | 3.8158 |  |  |  | 1.5 |  | 0.046417 | 2e-08
-1 | 10V | -0.99999 | 4.7 | [5 empty] | -1.3945 | 8e-08
-2 | 10V | -2 | 4.1 | [5 empty] | 0 | 4e-08
-5 | 10V | -5 | 4.1 | [5 empty] | -0.066666 | 2e-08
-10 | 10V | -10 | 3.8 |  |  |  | 0.5 |  | -0.060092 | 2e-08
-19 | 10V | -19.0001 | 3.8158 |  |  |  | 1.5 |  | -0.038236 | 1e-08
10 | 100V | 10.0003 | 3.8 | [5 empty] | 1.0929 | 1.1e-07
20 | 100V | 20.0003 | 3.8 | [5 empty] | 0.16666 | 5e-08
50 | 100V | 50.0005 | 7 | [5 empty] | 0.066666 | 3e-08
100 | 100V | 100.0008 | 6 |  |  |  | 2 |  | 0.066666 | 2e-08
-10 | 100V | -9.9998 | 3.8 | [5 empty] | -0.70712 | 1.1e-07
-20 | 100V | -19.9999 | 3.8 | [5 empty] | -0.25 | 5e-08
-50 | 100V | -50.0001 | 7 | [5 empty] | -0.1 | 3e-08
-100 | 100V | -100.0005 | 6 |  |  |  | 2 |  | -0.044096 | 2e-08
-100 | 1000V | -100.0007 | 6 | [5 empty] | -0.44096 | 4e-08
-200 | 1000V | -200.0012 | 5.5 | [5 empty] | -0.26352 | 2e-08
-500 | 1000V | -500.0045 | 8.2 | [5 empty] | -0.066666 | 2e-08
-1000 | 1000V | -1000.0092 | 7.6 |  |  |  | 2 |  | -0.05 | 1e-08
100 | 1000V | 100.0007 | 6 | [5 empty] | 0.7071 | 6e-08
200 | 1000V | 200.0014 | 5.5 | [5 empty] | 0 | 3e-08
500 | 1000V | 500.0045 | 8.2 | [5 empty] | 0.099999 | 2e-08
1000 | 1000V | 1000.0092 | 7.6 |  |  |  | 2 |  | 0.033333 | 1e-08
OHM PERFORMANCE TEST
OHM Test 4W | Calibrator ohm | Measured ohm | Source Unc. ppm | Lower Limit ohm | Upper Limit ohm | Measured ppm | DUT Transfer STB ppm | Test Result % of SPEC | MEAS SDEV ppm
1 Ω | 0.99961 | 0.9996 | 85 |  |  |  | 20 |  | 1.6421 | 1.55e-06
1.9 Ω | 1.9 | 1.8999 | 85 |  |  |  | 15 |  | 0.69628 | 8.3e-07
10 Ω | 10.0004 | 10.0003 | 26 |  |  |  | 5 |  | 0.13228 | 1.5e-07
19 Ω | 19.0039 | 19.0039 | 24 |  |  |  | 5 |  | 0.079417 | 9e-08
100 Ω | 100.0053 | 100.0056 | 15 |  |  |  | 3 |  | 0.049997 | 3e-08
190 Ω | 190.0032 | 190.0037 | 15 |  |  |  | 3 |  | 0.027739 | 2e-08
1 kΩ | 999.9469 | 999.9484 | 11 |  |  |  | 3 |  | 0.052707 | 3e-08
1.9 kΩ | 1899.9106 | 1899.9094 | 11 |  |  |  | 3 |  | 0.023209 | 2e-08
10 kΩ | 10000.453 | 10000.448 | 9 |  |  |  | 3 |  | 0 | 3e-08
19 kΩ | 18999.722 | 18999.669 | 9 |  |  |  | 3 |  | 0.017544 | 3e-08
100 kΩ | 99998.75 | 99998.34 | 11 |  |  |  | 5 |  | 0.052706 | 2e-08
190 kΩ | 189994.66 | 189993.76 | 11 |  |  |  | 5 |  | 0.070178 | 2e-08
1 MΩ | 999908.7 | 999902.6 | 16 |  |  |  | 8 |  | 0.15367 | 3e-08
1.9 MΩ | 1899855.1 | 1899831.1 | 17 |  |  |  | 8 |  | 0.017545 | 3e-08
10MΩ | 9998564 | 9998450 | 33 |  |  |  | 12 |  | 0.13231 | 1.3e-07
19 MΩ | 18999500 | 19000122 | 43 |  |  |  | 12 |  | 0.37029 | 6e-08
100MΩ(2W) | 100009390 | 100013990 | 110 |  |  |  | 180 |  | 1.4957 | 2.2e-07
OHM Zero 4W | DUT Range | Measured ohm | Source Unc. ohm | Lower Limit ohm | Upper Limit ohm | Measured | DUT Transfer STB ppm | Test Result % of SPEC | MEAS SDEV ohm
0 Ω | 10 Ω | -0.0001 | 5e-05 |  |  | N/A |  |  | 1.1559e-05 | 1.766e-06
0 Ω | 100 Ω | 0.00049 | 5e-05 |  |  | N/A |  |  | 5.2705e-06 | 2.31e-06
0 Ω | 1 KΩ | 0.005 | 5e-05 |  |  | N/A |  |  | 3.3333e-05 | 2.45e-05
0 Ω | 10 KΩ | 0.05 | 5e-05 |  |  | N/A |  |  | 0 | 0.000251
0 Ω | 100 KΩ | 0.14 | 5e-05 |  |  | N/A |  |  | 0 | 0.000672
0 Ω | 1 MΩ | 1.4 | 5e-05 |  |  | N/A |  |  | 0 | 0.00615
0 Ω | 10 MΩ | 14 | 5e-05 |  |  | N/A |  |  | 0 | 0.0599
0 Ω | 100 MΩ(2W) | 140 | 5e-05 |  |  | N/A |  |  | 0 | 0.567
OHM Zero 2W | DUT Range | Measured ohm | Source Unc. ohm | Lower Limit ohm | Upper Limit ohm | Measured | DUT Transfer STB ppm | Test Result % of SPEC | MEAS SDEV ohm
0 Ω | 10 Ω | 0.37616 | 5e-05 |  |  | N/A |  |  | 0.00012879 | 1.828e-06
0 Ω | 100 Ω | 0.37279 | 5e-05 |  |  | N/A |  |  | 4.5583e-05 | 2.78e-06
0 Ω | 1 KΩ | 0.3773 | 5e-05 |  |  | N/A |  |  | 8.8192e-05 | 2.4e-05
0 Ω | 10 KΩ | 0.424 | 5e-05 |  |  | N/A |  |  | 0.0005 | 0.000285
0 Ω | 100 KΩ | 0.52 | 5e-05 |  |  | N/A |  |  | 0 | 0.000729
0 Ω | 1 MΩ | 1.8 | 5e-05 |  |  | N/A |  |  | 2.3551e-16 | 0.00632
0 Ω | 10 MΩ | 15 | 5e-05 |  |  | N/A |  |  | 0 | 0.0658
0 Ω | 100 MΩ | 140 | 5e-05 |  |  | N/A |  |  | 0 | 0.595
ACV PERFORMANCE TEST
ACV Linearity Checks | DUT Range | Measured Vac | Source Unc. ppm | Lower Limit Vac | Upper Limit Vac | Measured ppm of Input | DUT Transfer STB ppm | Test Result % of SPEC | MEAS SDEV ppm
1VAC/1kHz | 10 VAC | 0.99912 | 77 |  |  |  | 10 |  | 68.0264 | 3.03e-06
2VAC/1kHz | 10 VAC | 1.9995 | 73.5 |  |  |  | 10 |  | 8.5817 | 1.46e-06
5VAC/1kHz | 10 VAC | 4.9998 | 84 |  |  |  | 10 |  | 1.3334 | 1.15e-06
10VAC/1kHz | 10 VAC | 10 | 77 |  |  |  | 10 |  | 1.1667 | 8.3e-07
12VAC/1kHz | 10 VAC | 12.0001 | 75.833 |  |  |  | 10 |  | 1.559 | 9.9e-07
15VAC/1kHz | 10 VAC | 15.0001 | 74.667 |  |  |  | 10 |  | 0.57734 | 9.7e-07
19VAC/1kHz | 10 VAC | 19.0003 | 73.684 |  |  |  | 10 |  | 0.91159 | 1.45e-06
ACV Performance Tset |  | Measured Vac | Source Unc. ppm | Lower Limit Vac | Upper Limit Vac | Measured ppm | DUT Transfer STB ppm | Test Result % of SPEC | MEAS SDEV ppm
0.001 V, 10 Hz |  | 0.0009959 | 5500 |  |  |  | 20 |  | 59.1918 | 5.89e-06
0.001 V, 20 Hz |  | 0.00099606 | 5200 |  |  |  | 20 |  | 15.8739 | 3.98e-06
0.001 V, 30 Hz |  | 0.00099611 | 5200 |  |  |  | 20 |  | 7.0987 | 3.27e-06
0.001 V, 40 Hz |  | 0.00099609 | 5100 |  |  |  | 20 |  | 7.848 | 5.53e-06
0.001 V, 55 Hz |  | 0.0009964 | 5100 |  |  |  | 20 |  | 8.3634 | 3.62e-06
0.001 V, 300 Hz |  | 0.00099637 | 5100 |  |  |  | 20 |  | 26.5539 | 2.44e-05
0.001 V, 1 KHz |  | 0.00099645 | 5100 |  |  |  | 20 |  | 17.4625 | 4.39e-06
0.001 V, 10 KHz |  | 0.00099647 | 5100 |  |  |  | 20 |  | 8.6909 | 5e-06
0.001 V, 20 KHz |  | 0.00099696 | 5100 |  |  |  | 20 |  | 7.287 | 7.76e-06
0.001 V, 30 KHz |  | 0.00099688 | 5340 |  |  |  | 20 |  | 6.0281 | 5.05e-06
0.001 V, 50 KHz |  | 0.00099727 | 5340 |  |  |  | 20 |  | 10.5698 | 2.95e-06
0.001 V, 100 KHz |  | 0.0010001 | 8800 |  |  |  | 50 |  | 0 | 5.18e-06
0.001 V, 200 KHz |  | 0.00099734 | 16100 |  |  |  | 100 |  | 14.8532 | 6.02e-06
0.001 V, 300 KHz |  | 0.00099467 | 16100 |  |  |  | 100 |  | 7.8592 | 4.89e-06
0.001 V, 500 KHz |  | 0.0010028 | 31500 |  |  |  | 200 |  | 17.9778 | 4.88e-05
0.001 V, 1 MHz |  | 0.0010105 | 33000 |  |  |  | 300 |  | 9.8958 | 1.378e-05
0.01 V, 10 Hz |  | 0.0099976 | 1100 |  |  |  | 20 |  | 5.1759 | 2.84e-06
0.01 V, 20 Hz |  | 0.009997 | 800 |  |  |  | 20 |  | 0.86629 | 2.47e-06
0.01 V, 30 Hz |  | 0.0099967 | 800 |  |  |  | 20 |  | 0.88221 | 5.6e-07
0.01 V, 40 Hz |  | 0.0099965 | 700 |  |  |  | 20 |  | 0.72674 | 2.55e-06
0.01 V, 55 Hz |  | 0.0099965 | 700 |  |  |  | 20 |  | 1.0142 | 2.66e-06
0.01 V, 300 Hz |  | 0.0099964 | 700 |  |  |  | 20 |  | 3.2714 | 2.35e-06
0.01 V, 1 KHz |  | 0.0099965 | 700 |  |  |  | 20 |  | 1.4245 | 1.8e-06
0.01 V, 10 KHz |  | 0.0099976 | 700 |  |  |  | 20 |  | 0.92818 | 2.5e-07
0.01 V, 20 KHz |  | 0.0099986 | 700 |  |  |  | 20 |  | 0.88204 | 8.3e-07
0.01 V, 30 KHz |  | 0.0099989 | 940 |  |  |  | 20 |  | 0.70718 | 6.3e-07
0.01 V, 50 KHz |  | 0.01 | 940 |  |  |  | 20 |  | 0.88189 | 4.7e-07
0.01 V, 100 KHz |  | 0.010005 | 1600 |  |  |  | 50 |  | 0.72613 | 6.5e-07
0.01 V, 200 KHz |  | 0.0099974 | 2600 |  |  |  | 100 |  | 1.8032 | 1.01e-06
0.01 V, 300 KHz |  | 0.0099946 | 2600 |  |  |  | 100 |  | 0.50027 | 4.7e-07
0.01 V, 500 KHz |  | 0.010006 | 4500 |  |  |  | 200 |  | 2.0603 | 4.7e-06
0.01 V, 1 MHz |  | 0.010057 | 6000 |  |  |  | 300 |  | 1.229 | 1.93e-06
0.1 V, 10 Hz |  | 0.10002 | 660 |  |  |  | 20 |  | 10.8355 | 9.71e-06
0.1 V, 20 Hz |  | 0.1 | 300 |  |  |  | 20 |  | 0.78171 | 2.02e-06
0.1 V, 30 Hz |  | 0.1 | 300 |  |  |  | 20 |  | 0.66665 | 2.09e-06
0.1 V, 40 Hz |  | 0.1 | 195 |  |  |  | 20 |  | 0.72647 | 1.51e-06
0.1 V, 55 Hz |  | 0.1 | 195 |  |  |  | 20 |  | 2.5055 | 2.28e-06
0.1 V, 300 Hz |  | 0.1 | 195 |  |  |  | 20 |  | 2.048 | 2.13e-06
0.1 V, 1 KHz |  | 0.1 | 195 |  |  |  | 20 |  | 1.236 | 1.01e-06
0.1 V, 10 KHz |  | 0.1 | 195 |  |  |  | 20 |  | 0.60092 | 3.7e-07
0.1 V, 20 KHz |  | 0.1 | 195 |  |  |  | 20 |  | 0.60092 | 6.1e-07
0.1 V, 30 KHz |  | 0.1 | 400 |  |  |  | 20 |  | 0.86602 | 6.9e-07
0.1 V, 50 KHz |  | 0.1 | 400 |  |  |  | 20 |  | 0.97183 | 8.4e-07
0.1 V, 100 KHz |  | 0.1 | 1050 |  |  |  | 50 |  | 1.0929 | 7.7e-07
0.1 V, 200 KHz |  | 0.10002 | 1240 |  |  |  | 100 |  | 0.60081 | 5.2e-07
0.1 V, 300 KHz |  | 0.099995 | 1240 |  |  |  | 100 |  | 0.60095 | 5.7e-07
0.1 V, 500 KHz |  | 0.099963 | 1900 |  |  |  | 200 |  | 0.92831 | 7.3e-07
0.1 V, 1 MHz |  | 0.10007 | 4000 |  |  |  | 300 |  | 1.832 | 1.27e-06
1 V, 10 Hz |  | 1.0001 | 600 |  |  |  | 10 |  | 9.9488 | 9.99e-06
1 V, 20 Hz |  | 1 | 180 |  |  |  | 10 |  | 0.8819 | 2.03e-06
1 V, 30 Hz |  | 1 | 180 |  |  |  | 10 |  | 0.66667 | 1.37e-06
1 V, 40 Hz |  | 0.99999 | 77 |  |  |  | 10 |  | 0.78175 | 2.43e-06
1 V, 55 Hz |  | 0.99999 | 77 |  |  |  | 10 |  | 3.9052 | 7.8e-07
1 V, 300 Hz |  | 0.99999 | 77 |  |  |  | 10 |  | 1.5 | 2.09e-06
1 V, 1 KHz |  | 1 | 77 |  |  |  | 10 |  | 1.4142 | 1.2e-06
1 V, 10 KHz |  | 0.99999 | 77 |  |  |  | 10 |  | 0.60093 | 4.5e-07
1 V, 20 KHz |  | 1 | 77 |  |  |  | 10 |  | 1.3944 | 4.7e-07
1 V, 30 KHz |  | 1 | 140 |  |  |  | 10 |  | 0.60091 | 2.3e-07
1 V, 50 KHz |  | 1.0001 | 140 |  |  |  | 20 |  | 0.7817 | 5.3e-07
1 V, 100 KHz |  | 0.99981 | 310 |  |  |  | 30 |  | 1.0002 | 8.2e-07
1 V, 200 KHz |  | 0.99938 | 550 |  |  |  | 70 |  | 0.50031 | 4.6e-07
1 V, 300 KHz |  | 0.99967 | 550 |  |  |  | 70 |  | 0.72672 | 3.5e-07
1 V, 500 KHz |  | 0.99901 | 1400 |  |  |  | 100 |  | 0.33366 | 2.9e-07
1 V, 1 MHz |  | 0.99972 | 3000 |  |  |  | 200 |  | 0.86627 | 1.48e-06
10 V, 10 Hz |  | 10.001 | 600 |  |  |  | 10 |  | 10.0709 | 6.9e-06
10 V, 20 Hz |  | 10 | 180 |  |  |  | 10 |  | 0.86602 | 2.53e-06
10 V, 30 Hz |  | 9.9999 | 180 |  |  |  | 10 |  | 1 | 2.08e-06
10 V, 40 Hz |  | 9.9998 | 77 |  |  |  | 10 |  | 0.86604 | 1.3e-06
10 V, 55 Hz |  | 9.9998 | 77 |  |  |  | 10 |  | 4.0621 | 2.41e-06
10 V, 300 Hz |  | 9.9998 | 77 |  |  |  | 10 |  | 1 | 1.9e-06
10 V, 1 KHz |  | 10 | 77 |  |  |  | 10 |  | 1.1667 | 9.2e-07
10 V, 10 KHz |  | 9.9998 | 77 |  |  |  | 10 |  | 0.44097 | 2.9e-07
10 V, 20 KHz |  | 10 | 77 |  |  |  | 10 |  | 1.5 | 6.1e-07
10 V, 30 KHz |  | 10.0001 | 140 |  |  |  | 10 |  | 0.49999 | 7e-07
10 V, 50 KHz |  | 10.0003 | 140 |  |  |  | 20 |  | 0.52703 | 2.9e-07
10 V, 100 KHz |  | 9.9981 | 270 |  |  |  | 30 |  | 0.88209 | 3.8e-07
10 V, 200 KHz |  | 9.9933 | 670 |  |  |  | 70 |  | 1.5645 | 5.5e-07
10 V, 300 KHz |  | 9.9962 | 670 |  |  |  | 70 |  | 2.439 | 5.2e-07
10 V, 500 KHz |  | 9.9897 | 1700 |  |  |  | 100 |  | 2.59 | 6.9e-07
10 V, 1 MHz |  | 10.005 | 3500 |  |  |  | 200 |  | 1.7392 | 2.52e-06
19 V, 10 Hz |  | 19.0067 | 552.63 |  |  |  | 10 |  | 6.0904 | 7.73e-06
19 V, 20 Hz |  | 19.0016 | 165.79 |  |  |  | 10 |  | 1.209 | 2.33e-06
19 V, 30 Hz |  | 19.0007 | 165.79 |  |  |  | 10 |  | 0.83215 | 1.06e-06
19 V, 40 Hz |  | 19.0003 | 73.684 |  |  |  | 10 |  | 0.83216 | 1.61e-06
19 V, 55 Hz |  | 19.0001 | 73.684 |  |  |  | 10 |  | 2.1016 | 1.69e-06
19 V, 300 Hz |  | 18.9998 | 73.684 |  |  |  | 10 |  | 2.4546 | 2.01e-06
19 V, 1 KHz |  | 18.9999 | 73.684 |  |  |  | 10 |  | 0.6446 | 1.49e-06
19 V, 10 KHz |  | 18.9998 | 73.684 |  |  |  | 10 |  | 0.61404 | 3.1e-07
19 V, 20 KHz |  | 19.0003 | 73.684 |  |  |  | 10 |  | 1.7009 | 4.7e-07
19 V, 30 KHz |  | 19.0006 | 130.53 |  |  |  | 10 |  | 0.63253 | 5.3e-07
19 V, 50 KHz |  | 19.0009 | 130.53 |  |  |  | 20 |  | 0.70172 | 2e-07
19 V, 100 KHz |  | 18.9969 | 251.05 |  |  |  | 30 |  | 0.77979 | 4.9e-07
19 V, 200 KHz |  | 18.9881 | 589.47 |  |  |  | 70 |  | 0.57557 | 6.6e-07
19 V, 300 KHz |  | 18.994 | 589.47 |  |  |  | 70 |  | 2.3397 | 3.7e-07
19 V, 500 KHz |  | 18.9819 | 1463.2 |  |  |  | 100 |  | 0.74503 | 3.5e-07
19 V, 1 MHz |  | 19.0142 | 3073.7 |  |  |  | 200 |  | 3.0866 | 8.4e-07
100 V, 10 Hz |  | 100.0112 | 600 |  |  |  | 10 |  | 8.5284 | 1.007e-05
100 V, 20 Hz |  | 100.0007 | 180 |  |  |  | 10 |  | 1.8708 | 2.79e-06
100 V, 30 Hz |  | 99.9993 | 180 |  |  |  | 10 |  | 1.118 | 1.36e-06
100 V, 40 Hz |  | 99.9982 | 85 |  |  |  | 10 |  | 0.88193 | 2.07e-06
100 V, 55 Hz |  | 99.9983 | 85 |  |  |  | 10 |  | 3.4076 | 1.01e-06
100 V, 300 Hz |  | 99.9976 | 85 |  |  |  | 10 |  | 1.9003 | 2.47e-06
100 V, 1 KHz |  | 99.9971 | 85 |  |  |  | 10 |  | 1.2248 | 1.17e-06
100 V, 10 KHz |  | 99.9969 | 85 |  |  |  | 10 |  | 0.52706 | 7.3e-07
100 V, 20 KHz |  | 99.9969 | 85 |  |  |  | 10 |  | 0.50002 | 2.2e-07
100 V, 30 KHz |  | 99.9978 | 240 |  |  |  | 10 |  | 0.60094 | 4.2e-07
100 V, 50 KHz |  | 99.997 | 240 |  |  |  | 20 |  | 1.2019 | 5.9e-07
100 V, 100 KHz |  | 99.9928 | 600 |  |  |  | 30 |  | 1.5001 | 8.5e-07
100 V, 200 KHz |  | 99.9958 | 2600 |  |  |  | 50 |  | 3.8407 | 3.48e-06
1000 V, 55 Hz |  | 1000.022 | 79 |  |  |  | 15 |  | 2.598 | 1.15e-06
1000 V, 300 Hz |  | 1000.033 | 79 |  |  |  | 15 |  | 2.4494 | 1.62e-06
700 V, 1 KHz |  | 700.004 | 80.714 |  |  |  | 50 |  | 1.2599 | 1.21e-06
DCI PERFORMANCE TEST
DCI Test | Expected Value | Measured A | Source Unc. ppm | Lower Limit A | Upper Limit A | Measured ppm | DUT Transfer STB ppm | Test Result % of SPEC | MEAS SDEV ppm
1 µADC | 1e-06 | 9.991e-07 | 10054.0486 | [5 empty] | 50.045 | 1.403e-05
10 µADC | 1e-05 | 9.9995e-06 | 1045.047 | [5 empty] | 5.7738 | 1.48e-06
50 µADC | 5e-05 | 5.0002e-05 | 245 | [5 empty] | 0 | 2e-07
100 µADC | 0.0001 | 0.0001 | 145 |  |  |  | 7 |  | 0.49998 | 1.1e-07
-100 µADC | -0.0001 | -0.00010001 | 145 |  |  |  | 7 |  | -0.49997 | 1.5e-07
-50 µADC | -5e-05 | -5.0004e-05 | 245 | [5 empty] | 0 | 2.7e-07
-10 µADC | -1e-05 | -1.0002e-05 | -1044.8328 | [5 empty] | 0 | 1.5e-06
0.1 mADC | 0.0001 | 9.9998e-05 | 145 | [5 empty] | 5.7736 | 1.05e-06
0.5 mADC | 0.0005 | 0.00050001 | 65 | [5 empty] | 0.99997 | 2.2e-07
1 mADC | 0.001 | 0.001 | 55 |  |  |  | 7 |  | 0 | 1.2e-07
-1 mADC | -0.001 | -0.001 | 55 |  |  |  | 7 |  | -0.57733 | 1.1e-07
-0.5 mADC | -0.0005 | -0.00050003 | 65 | [5 empty] | -0.99995 | 2.4e-07
1 mADC | 0.001 | 0.001 | 55 | [5 empty] | 4.9999 | 1.31e-06
5 mADC | 0.005 | 0.0050001 | 65 | [5 empty] | 0 | 2e-07
10 mADC | 0.01 | 0.01 | 55 |  |  |  | 7 |  | 0.49999 | 9e-08
-10 mADC | -0.01 | -0.01 | 55 |  |  |  | 7 |  | 0 | 1.3e-07
-5 mADC | -0.005 | -0.005 | 65 | [5 empty] | -1.1547 | 1.8e-07
10 mADC | 0.01 | 0.0099986 | 55 | [5 empty] | 0 | 1.17e-06
50 mADC | 0.05 | 0.05 | 75 | [5 empty] | 0 | 2.4e-07
100 mADC | 0.1 | 0.1 | 65 |  |  |  | 7 |  | 0.5 | 1e-07
-100 mADC | -0.1 | -0.1 | 65 |  |  |  | 7 |  | -0.49998 | 1.2e-07
-50 mADC | -0.05 | -0.050002 | 75 | [5 empty] | 0 | 2.3e-07
0.1 ADC | 0.1 | 0.1 | 65 | [5 empty] | 5.7732 | 1.34e-06
0.5 ADC | 0.5 | 0.50003 | 135 | [5 empty] | 1.1546 | 2.8e-07
-0.5 ADC | -0.5 | -0.50003 | 135 | [5 empty] | 0 | 2.3e-07
-1.0 ADC | -1 | -1 | 105 |  |  |  | 20 |  | 0 | 8e-08
1.0 ADC | 1 | 1 | 105 |  |  |  | 20 |  | 0.49998 | 1.7e-07
ACI PERFORMANCE TEST
ACI Test | Expected Value | Measured Aac | Source Unc. ppm | Lower Limit Aac | Upper Limit Aac | Measured ppm | DUT Transfer STB ppm | Test Result % of SPEC | MEAS SDEV ppm
10 µA, 10 Hz | 1e-05 | 1.0001e-05 | 3649.5438 | [5 empty] | 12.9083 | 1.345e-05
10 µA, 20 Hz | 1e-05 | 1.0001e-05 | 2849.8005 | [5 empty] | 4.9997 | 3.02e-06
10 µA, 30 Hz | 1e-05 | 1.0001e-05 | 2849.8005 | [5 empty] | 8.1644 | 2.46e-06
10 µA, 40 Hz | 1e-05 | 1e-05 | 2119.9046 | [5 empty] | 12.9094 | 7.15e-06
10 µA 55 Hz | 1e-05 | 9.9975e-06 | 2120.5301 | [5 empty] | 17.3248 | 1.718e-05
10 µA, 300 Hz | 1e-05 | 1e-05 | 2119.9152 |  |  | -40 |  |  | 12.5826 | 4.61e-06
10 µA, 1 KHz | 1e-05 | 9.9999e-06 | 2120.0318 | [5 empty] | 27.5383 | 2.68e-05
10 µA, 5 KHz | 1e-05 | 9.9968e-06 | 5501.7881 | [5 empty] | 28.7322 | 4.21e-06
10 µA, 10 KHz | 1e-05 | 9.9949e-06 | 11505.868 | [5 empty] | 0 | 7.54e-06
100 µA, 10 Hz | 0.0001 | 0.00010002 | 950 | 9.9939e-05 | 0.00010006 | -40 | 50 |  | 5.9148 | 9.02e-06
100 µA, 20 Hz | 0.0001 | 0.00010001 | 600 |  |  |  | 50 |  | 1.2582 | 2.03e-06
100 µA, 30 Hz | 0.0001 | 0.00010001 | 600 | 9.9939e-05 | 0.00010006 |  | 50 |  | 1.2909 | 8.2e-07
100 µA, 40 Hz | 0.0001 | 0.00010001 | 320 |  |  |  | 50 |  | 0.49996 | 1.92e-06
100 µA 55 Hz | 0.0001 | 0.00010001 | 320 | 9.9939e-05 | 0.00010006 | -40 | 50 |  | 4.349 | 8.3e-07
100 µA, 300 Hz | 0.0001 | 0.00010001 | 320 |  |  |  | 50 |  | 1.2582 | 1.95e-06
100 µA, 1 KHz | 0.0001 | 0.0001 | 320 | 9.9939e-05 | 0.00010006 |  | 50 |  | 2.2173 | 3.1e-06
100 µA, 5 KHz | 0.0001 | 0.0001 | 0.1 |  |  |  | 100 |  | 0.49998 | 6.9e-07
100 µA, 10 KHz | 0.0001 | 0.00010002 | 0.25 | 9.9939e-05 | 0.00010006 |  | 300 |  | 0.49989 | 9.6e-07
1mA, 10 Hz | 0.001 | 0.0010002 | 700 |  |  |  | 40 |  | 4.2024 | 5.7e-06
1 mA, 20 Hz | 0.001 | 0.0010001 | 390 | 9.9939e-05 | 0.00010006 |  | 40 |  | 2.0815 | 1.83e-06
1 mA, 30 Hz | 0.001 | 0.0010001 | 390 |  |  |  | 40 |  | 0.81644 | 2.31e-06
1 mA, 40 Hz | 0.001 | 0.001 | 160 | 9.9939e-05 | 0.00010006 |  | 40 |  | 0.49998 | 2.82e-06
1 mA, 55 Hz | 0.001 | 0.0010001 | 160 |  |  |  | 40 |  | 2.4999 | 2.19e-06
1 mA, 300 Hz | 0.001 | 0.001 | 160 | 9.9939e-05 | 0.00010006 |  | 40 |  | 2.0615 | 1.93e-06
1 mA, 1 KHz | 0.001 | 0.001 | 160 |  |  |  | 40 |  | 1.2909 | 1.1e-06
1 mA, 5 KHz | 0.001 | 0.001 | 1000 | 9.9939e-05 | 0.00010006 |  | 70 |  | 0.99997 | 3.7e-07
1 mA, 10 KHz | 0.001 | 0.001 | 0.25 |  |  |  | 200 |  | 0.57735 | 3.8e-07
10 mA, 10 Hz | 0.01 | 0.010002 | 700 | 9.9939e-05 | 0.00010006 |  | 40 |  | 6.8483 | 1.034e-05
10 mA, 20 Hz | 0.01 | 0.010001 | 390 |  |  |  | 40 |  | 0.99991 | 2.19e-06
10 mA, 30 Hz | 0.01 | 0.010001 | 390 | 9.9939e-05 | 0.00010006 |  | 40 |  | 0.49997 | 1.51e-06
10 mA, 40 Hz | 0.01 | 0.010001 | 160 |  |  |  | 40 |  | 0.95737 | 1.85e-06
10 mA, 55 Hz | 0.01 | 0.01 | 160 | 9.9939e-05 | 0.00010006 |  | 40 |  | 2.8721 | 6.5e-07
10 mA, 300 Hz | 0.01 | 0.010001 | 160 |  |  |  | 40 |  | 1.4141 | 2.04e-06
10 mA, 1 KHz | 0.01 | 0.010001 | 160 | 9.9939e-05 | 0.00010006 |  | 40 |  | 0.95735 | 1.35e-06
10 mA, 5 KHz | 0.01 | 0.010002 | 1000 |  |  |  | 70 |  | 0.95727 | 9.9e-07
10 mA, 10 KHz | 0.01 | 0.01 | 0.25 | 9.9939e-05 | 0.00010006 |  | 200 |  | 0.81649 | 6.1e-07
100 mA, 10 Hz | 0.1 | 0.10002 | 700 |  |  |  | 40 |  | 4.271 | 7.6e-06
100 mA, 20 Hz | 0.1 | 0.10001 | 390 | 9.9939e-05 | 0.00010006 |  | 40 |  | 2.8864 | 2.27e-06
100 mA, 30 Hz | 0.1 | 0.10001 | 390 |  |  |  | 40 |  | 0.8164 | 2.45e-06
100 mA, 40 Hz | 0.1 | 0.10001 | 160 | 9.9939e-05 | 0.00010006 |  | 40 |  | 0.49995 | 2.26e-06
100 mA, 55 Hz | 0.1 | 0.10001 | 160 |  |  |  | 40 |  | 3.4153 | 1.75e-06
100 mA, 300 Hz | 0.1 | 0.10001 | 160 | 9.9939e-05 | 0.00010006 |  | 40 |  | 2.7077 | 1.38e-06
100 mA, 1 KHz | 0.1 | 0.10001 | 160 |  |  |  | 40 |  | 1.1546 | 7e-07
100 mA, 5 KHz | 0.1 | 0.1 | 0.1 | 9.9939e-05 | 0.00010006 |  | 70 |  | 1.2583 | 5e-07
100 mA, 10 KHz | 0.1 | 0.099995 | 0.25 |  |  |  | 200 |  | 0.81654 | 3.1e-07
1 A, 10 Hz | 1 | 1 | -1 | 9.9939e-05 | 0.00010006 |  | 40 |  | 8.1392 | 1.087e-05
1 A, 20 Hz | 1 | 1 | -1 |  |  |  | 40 |  | 1.7078 | 2.74e-06
1 A, 30 Hz | 1 | 1 | -1 | 9.9939e-05 | 0.00010006 |  | 40 |  | 1.2583 | 1.06e-06
1 A, 40 Hz | 1 | 1.0001 | 640 |  |  |  | 40 |  | 0.81638 | 2.78e-06
1 A, 55 Hz | 1 | 1.0001 | 640 | 9.9939e-05 | 0.00010006 |  | 40 |  | 3.3162 | 1.73e-06
1 A, 300 Hz | 1 | 1.0001 | 640 |  |  |  | 40 |  | 0.9573 | 2.58e-06
1 A, 1 KHz | 1 | 1.0001 | 640 | 9.9939e-05 | 0.00010006 |  | 40 |  | 0.95733 | 1.16e-06
1 A, 5 KHz | 1 | 0.99946 | 800 |  |  |  | 70 |  | 0.57766 | 6.1e-07
1 A, 10 KHz | 1 | 0.99996 | 0.82 | 9.9939e-05 | 0.00010006 |  | 200 |  | 0.50002 | 3.3e-07
Test Summary | Initialization | DCV | OHM | ACV | DCI | ACI |  | Total
Time escape minute | 0.19 | 206.89 | 249.13 | 652.49 | 40.39 | 185.12