## Sheet
PERFORMANCE TEST
DUT
Manufacturer |  | Datron Instruments |  |  | S/N |  | 31255
Model Number |  | 4950 |  |  | Test date |  | 23-09-06 22:20:43
Meter Info |  | Datron 4950 STD |  |  | Ambient Temp. |  | 23.0
Options |  | 1,1,1 |  |  | CERT_AMB? |  | 23
Firmware |  | 890197-02.07 |  |  | TEMP? |  | 43.46
Date of Baseline Cal |  | "1999-SEP-16" |  |  | TEMP? BASE |  | 3.46
Date of Certified Cal |  | "2023-SEP-04" |  |  | TEMP? CERTIFIED |  | 0.07
CNFTST? |  | 0 |  |  | STANDBY? |  | "00:00"
DUT Confidence level |  | 30 Days Transfer Satbility
Calibrator
Reference standard |  | Options |  | Unc | Calibration Date |  | Due Date
Fluke 5700A |  |  |  | -- | -- |  | --
ADRmu12 |  | 10.000000 VDC |  | -- | -- |  | --
Fluke 742A-1 8927002 |  | 1.0000195 ohm |  | -- | -- |  | --
Fluke 742A-10k 7815006 |  | 10.000133 kohm
MFC
MFC Calibrate Date |  | 030622 |  |  | MFC last calibrated |  | 460 days
Calibrate Date WBFLAT |  | 100589 |  |  | MFC since WBFLAT |  | 0 days
Calibrate Date WB Gain |  | 250190 |  |  | MFC since WBGAIN |  | 0 days
MFC Calibrate Date Zero |  | 270823 |  |  | MFC since DCV Zero |  | 10 days
MFC Confidence level |  | 5700A 99% 90 Days |  |  | CAL CONST 11V Range +zero |  | 398.1791
CAL CONST 6.5V |  | +6.529921596265443E+00 |  |  | CAL CONST 11V Range -zero |  | 398.1783
CAL CONST 13V reference voltage |  | +1.306136884003510E+01 |  |  | CAL CONST, Zero calibration temperature |  | 23
CAL CONST 10KOHM true output resistance |  | +1.000039902274076E+04 |  |  | CAL CONST, ALL calibration temp |  | 23
CAL SONST 10kOHM standard resistance |  | +1.000059914344082E+04 |  |  | CAL_CONST? CHECK, WB_FLAT_TEMP |  | 23.9
CAL CONST 1 OHM |  | +9.996206897944441E-01 |  |  | CAL CONST,WBAND_TEMP
ETIME? |  | 6046 days 22.92 hours |  |  | CAL TEMP? CHECK
ONTIME? |  | 2 days 8.78 hours |  |  | CAL CONST,DACLIN |  | 0.359
DCV PERFORMANCE TEST
DCV Test VDC | Expected Value | Measured μV | Source Unc. μV | Lower Limit μV | Upper Limit μV | Deviation | DUT Transfer STB ppm | Test Result % of SPEC | MEAS SDEV μV
0.0mVDC | 0 | 0.03 | 7.5e-07 |  |  | N/A |  |  | 0.029486 | 0.00622
0.0VDC | 0 | -0.2 | 7.5e-07 |  |  | N/A |  |  | 0.086603 | 0.01236
0.0VDC | 0 | 2 | 7.5e-07 |  |  | N/A |  |  | 0 | 0.0767
0.0VDC | 0 | 20 | 7.5e-07 |  |  | N/A |  |  | 5 | 0.941
0.0VDC | 0 | 100 | 7.5e-07 |  |  | N/A |  |  | 50 | 3.97
DCV Test VDC | DMM Range | Measured V | Source Unc. ppm | Lower Limit V | Upper Limit V | Measured ppm | DUT Transfer STB ppm | Test Result % of SPEC | MEAS SDEV ppm
0.01 | 100mV | 0.01 | 81.5 | [5 empty] | 1.6158 | 5.6e-07
0.02 | 100mV | 0.02 | 44 | [5 empty] | 0.74998 | 2.8e-07
0.05 | 100mV | 0.05 | 21.5 | [5 empty] | 0.2 | 1.1e-07
0.1 | 100mV | 0.1 | 14 |  |  |  | 3 |  | 0.16667 | 7e-08
-0.01 | 100mV | -0.0099994 | -81.5 | [5 empty] | -0.86608 | 5.3e-07
-0.02 | 100mV | -0.019999 | -44 | [5 empty] | -0.22049 | 2.7e-07
-0.05 | 100mV | -0.049999 | 21.5 | [5 empty] | -0.19437 | 1.1e-07
-0.1 | 100mV | -0.099999 | 14 |  |  |  | 3 |  | -0.14814 | 7e-08
0.1 | 1V | 0.1 | 14 | [5 empty] | 0.78174 | 1.2e-07
0.2 | 1V | 0.2 | 10.25 | [5 empty] | 0.35355 | 9e-08
0.5 | 1V | 0.5 | 5.9 | [5 empty] | 0.14142 | 3e-08
1 | 1V | 1 | 4.7 |  |  |  | 1.5 |  | 0.13333 | 2e-08
-0.1 | 1V | -0.099999 | 14 | [5 empty] | -0.97184 | 1.4e-07
-0.2 | 1V | -0.2 | 10.25 | [5 empty] | -0.25 | 7e-08
-0.5 | 1V | -0.5 | 5.9 | [5 empty] | -0.1 | 3e-08
-1 | 1V | -1 | 4.7 |  |  |  | 1.5 |  | -0.060093 | 2e-08
1 | 10V | 1 | 4.7 | [5 empty] | 0.72648 | 7e-08
2 | 10V | 2 | 4.1 | [5 empty] | 0.33333 | 5e-08
5 | 10V | 5 | 4.1 | [5 empty] | 0.066667 | 3e-08
10 | 10V | 10 | 3.8 |  |  |  | 1.5 |  | 0.086603 | 2e-08
19 | 10V | 19 | 3.8158 |  |  |  | 1.5 |  | 0.052632 | 2e-08
-1 | 10V | -1 | 4.7 | [5 empty] | -1.6667 | 7e-08
-2 | 10V | -2 | 4.1 | [5 empty] | -0.25 | 6e-08
-5 | 10V | -5 | 4.1 | [5 empty] | -0.088192 | 2e-08
-10 | 10V | -10 | 3.8 |  |  |  | 0.5 |  | -0.060093 | 2e-08
-19 | 10V | -19 | 3.8158 |  |  |  | 1.5 |  | -0.076472 | 2e-08
10 | 100V | 10.0001 | 3.8 | [5 empty] | 1.4814 | 1.1e-07
20 | 100V | 20.0002 | 3.8 | [5 empty] | 0.25 | 6e-08
50 | 100V | 50.0004 | 7 | [5 empty] | 0.099999 | 3e-08
100 | 100V | 100.0007 | 6 |  |  |  | 2 |  | 0.052704 | 2e-08
-10 | 100V | -9.9999 | 3.8 | [5 empty] | -0.70712 | 8e-08
-20 | 100V | -19.9998 | 3.8 | [5 empty] | -0.25 | 5e-08
-50 | 100V | -49.9995 | 7 | [5 empty] | -0.088193 | 3e-08
-100 | 100V | -99.9991 | 6 |  |  |  | 2 |  | -0.078174 | 2e-08
-100 | 1000V | -99.9999 | 6 | [5 empty] | -0.5 | 4e-08
-200 | 1000V | -199.9997 | 5.5 | [5 empty] | -0.25 | 2e-08
-500 | 1000V | -500.0012 | 8.2 | [5 empty] | 0 | 2e-08
-1000 | 1000V | -1000.0026 | 7.6 |  |  |  | 2 |  | -0.078173 | 2e-08
100 | 1000V | 100 | 6 | [5 empty] | 1.2247 | 6e-08
200 | 1000V | 200 | 5.5 | [5 empty] | 0.22048 | 2e-08
500 | 1000V | 500.0012 | 8.2 | [5 empty] | 0.1 | 2e-08
1000 | 1000V | 1000.0026 | 7.6 |  |  |  | 2 |  | 0.10138 | 2e-08
OHM PERFORMANCE TEST
OHM Test 4W | Calibrator ohm | Measured ohm | Source Unc. ppm | Lower Limit ohm | Upper Limit ohm | Measured ppm | DUT Transfer STB ppm | Test Result % of SPEC | MEAS SDEV ppm
1 Ω | 0.99962 | 0.9996 | 85 |  |  |  | 20 |  | 1.7166 | 1.75e-06
1.9 Ω | 1.9 | 1.8999 | 85 |  |  |  | 15 |  | 0.58847 | 9.8e-07
10 Ω | 10.0003 | 10.0003 | 26 |  |  |  | 5 |  | 0.1118 | 1.6e-07
19 Ω | 19.0037 | 19.0038 | 24 |  |  |  | 5 |  | 0.098445 | 9e-08
100 Ω | 100.0049 | 100.0052 | 15 |  |  |  | 3 |  | 0.097177 | 3e-08
190 Ω | 190.0021 | 190.0031 | 15 |  |  |  | 3 |  | 0.023208 | 2e-08
1 kΩ | 999.9422 | 999.945 | 11 |  |  |  | 3 |  | 0.06667 | 3e-08
1.9 kΩ | 1899.9031 | 1899.9046 | 11 |  |  |  | 3 |  | 0.037218 | 2e-08
10 kΩ | 10000.399 | 10000.416 | 9 |  |  |  | 3 |  | 0.12692 | 4e-08
19 kΩ | 18999.631 | 18999.628 | 9 |  |  |  | 3 |  | 0.023209 | 2e-08
100 kΩ | 99998.19 | 99998.11 | 11 |  |  |  | 5 |  | 0.11667 | 2e-08
190 kΩ | 189993.6 | 189993.47 | 11 |  |  |  | 5 |  | 0.089028 | 2e-08
1 MΩ | 999904.9 | 999898.9 | 16 |  |  |  | 8 |  | 0.1871 | 5e-08
1.9 MΩ | 1899846.3 | 1899822.5 | 17 |  |  |  | 8 |  | 0.031631 | 5e-08
10MΩ | 9998563 | 9998273 | 33 |  |  |  | 12 |  | 0.086617 | 1e-07
19 MΩ | 18999517 | 18999420 | 43 |  |  |  | 12 |  | 0.23407 | 6e-08
100MΩ(2W) | 100007650 | 99956840 | 110 |  |  |  | 180 |  | 1.5771 | 2.1e-07
OHM Zero 4W | DUT Range | Measured ohm | Source Unc. ohm | Lower Limit ohm | Upper Limit ohm | Measured | DUT Transfer STB ppm | Test Result % of SPEC | MEAS SDEV ohm
0 Ω | 10 Ω | -0.000107 | 5e-05 |  |  | N/A |  |  | 1.1874e-05 | 1.544e-06
0 Ω | 100 Ω | 0.00016 | 5e-05 |  |  | N/A |  |  | 5e-06 | 2.49e-06
0 Ω | 1 KΩ | 0.0018 | 5e-05 |  |  | N/A |  |  | 6.0093e-05 | 2.39e-05
0 Ω | 10 KΩ | 0.018 | 5e-05 |  |  | N/A |  |  | 0.0005 | 0.000239
0 Ω | 100 KΩ | 0.01 | 5e-05 |  |  | N/A |  |  | 0.0044096 | 0.000739
0 Ω | 1 MΩ | 0.1 | 5e-05 |  |  | N/A |  |  | 0.060093 | 0.00609
0 Ω | 10 MΩ | 0 | 5e-05 |  |  | N/A |  |  | 0.44096 | 0.0573
0 Ω | 100 MΩ(2W) | 10 | 5e-05 |  |  | N/A |  |  | 5 | 0.563
OHM Zero 2W | DUT Range | Measured ohm | Source Unc. ohm | Lower Limit ohm | Upper Limit ohm | Measured | DUT Transfer STB ppm | Test Result % of SPEC | MEAS SDEV ohm
0 Ω | 10 Ω | 0.20598 | 5e-05 |  |  | N/A |  |  | 0.00023176 | 2.29e-06
0 Ω | 100 Ω | 0.20207 | 5e-05 |  |  | N/A |  |  | 8.5991e-05 | 2.67e-06
0 Ω | 1 KΩ | 0.2033 | 5e-05 |  |  | N/A |  |  | 6.6667e-05 | 2.66e-05
0 Ω | 10 KΩ | 0.219 | 5e-05 |  |  | N/A |  |  | 0.00070711 | 0.000218
0 Ω | 100 KΩ | 0.21 | 5e-05 |  |  | N/A |  |  | 0.0044096 | 0.000668
0 Ω | 1 MΩ | 0.3 | 5e-05 |  |  | N/A |  |  | 0.044096 | 0.00658
0 Ω | 10 MΩ | 1 | 5e-05 |  |  | N/A |  |  | 0.52705 | 0.0459
0 Ω | 100 MΩ | 10 | 5e-05 |  |  | N/A |  |  | 5 | 0.536
ACV PERFORMANCE TEST
ACV Linearity Checks | DUT Range | Measured Vac | Source Unc. ppm | Lower Limit Vac | Upper Limit Vac | Measured ppm of Input | DUT Transfer STB ppm | Test Result % of SPEC | MEAS SDEV ppm
1VAC/1kHz | 10 VAC | 0.99939 | 77 |  |  |  | 10 |  | 12.2549 | 1.91e-06
2VAC/1kHz | 10 VAC | 1.9996 | 73.5 |  |  |  | 10 |  | 2.6357 | 2e-06
5VAC/1kHz | 10 VAC | 4.9999 | 84 |  |  |  | 10 |  | 1.7639 | 6.8e-07
10VAC/1kHz | 10 VAC | 10.0001 | 77 |  |  |  | 10 |  | 0.92795 | 1.37e-06
12VAC/1kHz | 10 VAC | 12.0001 | 75.833 |  |  |  | 10 |  | 1.0577 | 1.17e-06
15VAC/1kHz | 10 VAC | 15.0002 | 74.667 |  |  |  | 10 |  | 1.0244 | 1.47e-06
19VAC/1kHz | 10 VAC | 19.0003 | 73.684 |  |  |  | 10 |  | 0.95286 | 9.1e-07
ACV Performance Tset |  | Measured Vac | Source Unc. ppm | Lower Limit Vac | Upper Limit Vac | Measured ppm | DUT Transfer STB ppm | Test Result % of SPEC | MEAS SDEV ppm
0.001 V, 10 Hz |  | 0.0009899 | 5500 |  |  |  | 20 |  | 56.9221 | 4.37e-05
0.001 V, 20 Hz |  | 0.00098967 | 5200 |  |  |  | 20 |  | 142.1912 | 3.9e-05
0.001 V, 30 Hz |  | 0.00099025 | 5200 |  |  |  | 20 |  | 223.7163 | 1.256e-05
0.001 V, 40 Hz |  | 0.00098939 | 5100 |  |  |  | 20 |  | 204.85 | 1.474e-05
0.001 V, 55 Hz |  | 0.00099035 | 5100 |  |  |  | 20 |  | 129.8785 | 7.3e-06
0.001 V, 300 Hz |  | 0.00099062 | 5100 |  |  |  | 20 |  | 67.9051 | 2.63e-05
0.001 V, 1 KHz |  | 0.00099059 | 5100 |  |  |  | 20 |  | 493.7783 | 3.75e-05
0.001 V, 10 KHz |  | 0.00099068 | 5100 |  |  |  | 20 |  | 74.7651 | 8.16e-06
0.001 V, 20 KHz |  | 0.00099055 | 5100 |  |  |  | 20 |  | 71.6426 | 9.13e-05
0.001 V, 30 KHz |  | 0.00099065 | 5340 |  |  |  | 20 |  | 27.232 | 1.817e-05
0.001 V, 50 KHz |  | 0.00099081 | 5340 |  |  |  | 20 |  | 35.3246 | 1.42e-05
0.001 V, 100 KHz |  | 0.00099441 | 8800 |  |  |  | 50 |  | 31.7121 | 3.1e-05
0.001 V, 200 KHz |  | 0.00099173 | 16100 |  |  |  | 100 |  | 33.8207 | 3.41e-05
0.001 V, 300 KHz |  | 0.00098877 | 16100 |  |  |  | 100 |  | 21.7174 | 2.8e-05
0.001 V, 500 KHz |  | 0.00099513 | 31500 |  |  |  | 200 |  | 18.1932 | 5.58e-05
0.001 V, 1 MHz |  | 0.00099005 | 33000 |  |  |  | 300 |  | 24.9691 | 2.34e-05
0.01 V, 10 Hz |  | 0.0099986 | 1100 |  |  |  | 20 |  | 1.1182 | 2.14e-06
0.01 V, 20 Hz |  | 0.0099978 | 800 |  |  |  | 20 |  | 6.9797 | 7.6e-07
0.01 V, 30 Hz |  | 0.0099974 | 800 |  |  |  | 20 |  | 0.88215 | 1.72e-06
0.01 V, 40 Hz |  | 0.0099972 | 700 |  |  |  | 20 |  | 4.3633 | 8.1e-07
0.01 V, 55 Hz |  | 0.0099972 | 700 |  |  |  | 20 |  | 1.2022 | 2.11e-06
0.01 V, 300 Hz |  | 0.0099972 | 700 |  |  |  | 20 |  | 2.7139 | 1.05e-06
0.01 V, 1 KHz |  | 0.0099973 | 700 |  |  |  | 20 |  | 1.8033 | 1.59e-06
0.01 V, 10 KHz |  | 0.0099984 | 700 |  |  |  | 20 |  | 0.78186 | 9.7e-07
0.01 V, 20 KHz |  | 0.0099992 | 700 |  |  |  | 20 |  | 10.5061 | 1.21e-06
0.01 V, 30 KHz |  | 0.0099997 | 940 |  |  |  | 20 |  | 1.0929 | 5.9e-07
0.01 V, 50 KHz |  | 0.010001 | 940 |  |  |  | 20 |  | 14.6396 | 4e-07
0.01 V, 100 KHz |  | 0.010005 | 1600 |  |  |  | 50 |  | 1.0535 | 8e-07
0.01 V, 200 KHz |  | 0.0099978 | 2600 |  |  |  | 100 |  | 1.6919 | 7.3e-07
0.01 V, 300 KHz |  | 0.0099948 | 2600 |  |  |  | 100 |  | 0.50026 | 8e-07
0.01 V, 500 KHz |  | 0.010005 | 4500 |  |  |  | 200 |  | 2.006 | 9.2e-07
0.01 V, 1 MHz |  | 0.010042 | 6000 |  |  |  | 300 |  | 4.1957 | 3.17e-06
0.1 V, 10 Hz |  | 0.10002 | 660 |  |  |  | 20 |  | 8.9428 | 7.6e-06
0.1 V, 20 Hz |  | 0.1 | 300 |  |  |  | 20 |  | 1.3642 | 2.16e-06
0.1 V, 30 Hz |  | 0.1 | 300 |  |  |  | 20 |  | 0.49998 | 1.54e-06
0.1 V, 40 Hz |  | 0.1 | 195 |  |  |  | 20 |  | 1.0138 | 1.47e-06
0.1 V, 55 Hz |  | 0.1 | 195 |  |  |  | 20 |  | 1.4529 | 2.24e-06
0.1 V, 300 Hz |  | 0.1 | 195 |  |  |  | 20 |  | 2.2973 | 1.68e-06
0.1 V, 1 KHz |  | 0.1 | 195 |  |  |  | 20 |  | 1.0929 | 1.22e-06
0.1 V, 10 KHz |  | 0.1 | 195 |  |  |  | 20 |  | 0.7071 | 4.1e-07
0.1 V, 20 KHz |  | 0.1 | 195 |  |  |  | 20 |  | 1.4142 | 5.6e-07
0.1 V, 30 KHz |  | 0.1 | 400 |  |  |  | 20 |  | 0.52704 | 4.9e-07
0.1 V, 50 KHz |  | 0.1 | 400 |  |  |  | 20 |  | 0.5 | 3.4e-07
0.1 V, 100 KHz |  | 0.1 | 1050 |  |  |  | 50 |  | 0.5 | 7.7e-07
0.1 V, 200 KHz |  | 0.10002 | 1240 |  |  |  | 100 |  | 0.44088 | 5.3e-07
0.1 V, 300 KHz |  | 0.099991 | 1240 |  |  |  | 100 |  | 0.8661 | 7.3e-07
0.1 V, 500 KHz |  | 0.099958 | 1900 |  |  |  | 200 |  | 0.44115 | 1.6e-07
0.1 V, 1 MHz |  | 0.10041 | 4000 |  |  |  | 300 |  | 1.3888 | 5.9e-07
1 V, 10 Hz |  | 1.0001 | 600 |  |  |  | 10 |  | 8.1607 | 8.99e-06
1 V, 20 Hz |  | 1 | 180 |  |  |  | 10 |  | 1.9003 | 1.5e-06
1 V, 30 Hz |  | 1 | 180 |  |  |  | 10 |  | 1 | 1.45e-06
1 V, 40 Hz |  | 0.99999 | 77 |  |  |  | 10 |  | 1.2019 | 7.6e-07
1 V, 55 Hz |  | 0.99999 | 77 |  |  |  | 10 |  | 0.70712 | 2.2e-06
1 V, 300 Hz |  | 0.99999 | 77 |  |  |  | 10 |  | 2.3979 | 1.68e-06
1 V, 1 KHz |  | 1 | 77 |  |  |  | 10 |  | 1.3229 | 1.02e-06
1 V, 10 KHz |  | 0.99999 | 77 |  |  |  | 10 |  | 0.52705 | 4.7e-07
1 V, 20 KHz |  | 1 | 77 |  |  |  | 10 |  | 0.52704 | 5.9e-07
1 V, 30 KHz |  | 1 | 140 |  |  |  | 10 |  | 0 | 3.2e-07
1 V, 50 KHz |  | 1 | 140 |  |  |  | 20 |  | 0.49998 | 2.9e-07
1 V, 100 KHz |  | 0.9998 | 310 |  |  |  | 30 |  | 1.014 | 5.2e-07
1 V, 200 KHz |  | 0.99937 | 550 |  |  |  | 70 |  | 0.70756 | 3.4e-07
1 V, 300 KHz |  | 0.99963 | 550 |  |  |  | 70 |  | 0.52724 | 4.4e-07
1 V, 500 KHz |  | 0.99896 | 1400 |  |  |  | 100 |  | 0.50052 | 4e-07
1 V, 1 MHz |  | 0.99586 | 3000 |  |  |  | 200 |  | 1.4002 | 9.9e-07
10 V, 10 Hz |  | 10.0011 | 600 |  |  |  | 10 |  | 8.0424 | 6.16e-06
10 V, 20 Hz |  | 10.0002 | 180 |  |  |  | 10 |  | 1.424 | 2.18e-06
10 V, 30 Hz |  | 10 | 180 |  |  |  | 10 |  | 1.1304 | 1.19e-06
10 V, 40 Hz |  | 9.9999 | 77 |  |  |  | 10 |  | 1.3642 | 8.6e-07
10 V, 55 Hz |  | 9.9999 | 77 |  |  |  | 10 |  | 1.2019 | 2.35e-06
10 V, 300 Hz |  | 9.9999 | 77 |  |  |  | 10 |  | 2.5222 | 9.8e-07
10 V, 1 KHz |  | 10 | 77 |  |  |  | 10 |  | 1.3642 | 1.09e-06
10 V, 10 KHz |  | 9.9999 | 77 |  |  |  | 10 |  | 0.44096 | 2.9e-07
10 V, 20 KHz |  | 10.0001 | 77 |  |  |  | 10 |  | 0.97181 | 5.5e-07
10 V, 30 KHz |  | 10.0002 | 140 |  |  |  | 10 |  | 0.52704 | 3.9e-07
10 V, 50 KHz |  | 10.0003 | 140 |  |  |  | 20 |  | 0.78171 | 4.1e-07
10 V, 100 KHz |  | 9.998 | 270 |  |  |  | 30 |  | 1.2363 | 3.1e-07
10 V, 200 KHz |  | 9.9931 | 670 |  |  |  | 70 |  | 0.78227 | 3e-07
10 V, 300 KHz |  | 9.9962 | 670 |  |  |  | 70 |  | 1.0142 | 6.2e-07
10 V, 500 KHz |  | 9.9894 | 1700 |  |  |  | 100 |  | 0.60156 | 5.5e-07
10 V, 1 MHz |  | 9.958 | 3500 |  |  |  | 200 |  | 1.7232 | 2.52e-06
19 V, 10 Hz |  | 19.0068 | 552.63 |  |  |  | 10 |  | 6.0129 | 6.27e-06
19 V, 20 Hz |  | 19.0018 | 165.79 |  |  |  | 10 |  | 1.301 | 2.2e-06
19 V, 30 Hz |  | 19.0009 | 165.79 |  |  |  | 10 |  | 0.63252 | 9.9e-07
19 V, 40 Hz |  | 19.0005 | 73.684 |  |  |  | 10 |  | 0.74946 | 1.32e-06
19 V, 55 Hz |  | 19.0002 | 73.684 |  |  |  | 10 |  | 1.6804 | 2.25e-06
19 V, 300 Hz |  | 19 | 73.684 |  |  |  | 10 |  | 3.0349 | 1.76e-06
19 V, 1 KHz |  | 19.0001 | 73.684 |  |  |  | 10 |  | 1.0779 | 1.27e-06
19 V, 10 KHz |  | 19 | 73.684 |  |  |  | 10 |  | 0.55479 | 3.2e-07
19 V, 20 KHz |  | 19.0005 | 73.684 |  |  |  | 10 |  | 0.7443 | 6.3e-07
19 V, 30 KHz |  | 19.0006 | 130.53 |  |  |  | 10 |  | 0.59492 | 7.7e-07
19 V, 50 KHz |  | 19.001 | 130.53 |  |  |  | 20 |  | 0.55476 | 6.8e-07
19 V, 100 KHz |  | 18.9968 | 251.05 |  |  |  | 30 |  | 0.91176 | 3.7e-07
19 V, 200 KHz |  | 18.9878 | 589.47 |  |  |  | 70 |  | 0.65096 | 6e-07
19 V, 300 KHz |  | 18.9936 | 589.47 |  |  |  | 70 |  | 0.87749 | 9.8e-07
19 V, 500 KHz |  | 18.9808 | 1463.2 |  |  |  | 100 |  | 0.53411 | 7.3e-07
19 V, 1 MHz |  | 18.9239 | 3073.7 |  |  |  | 200 |  | 1.4711 | 9.3e-07
100 V, 10 Hz |  | 100.0114 | 600 |  |  |  | 10 |  | 11.746 | 7.8e-06
100 V, 20 Hz |  | 100.0014 | 180 |  |  |  | 10 |  | 1.3017 | 1.92e-06
100 V, 30 Hz |  | 99.9998 | 180 |  |  |  | 10 |  | 1.1304 | 1.42e-06
100 V, 40 Hz |  | 99.9987 | 85 |  |  |  | 10 |  | 1.2248 | 1.11e-06
100 V, 55 Hz |  | 99.9988 | 85 |  |  |  | 10 |  | 1.3333 | 2.92e-06
100 V, 300 Hz |  | 99.9981 | 85 |  |  |  | 10 |  | 1.5812 | 1.84e-06
100 V, 1 KHz |  | 99.9976 | 85 |  |  |  | 10 |  | 1.0929 | 1.34e-06
100 V, 10 KHz |  | 99.9979 | 85 |  |  |  | 10 |  | 0.78175 | 4.9e-07
100 V, 20 KHz |  | 99.9973 | 85 |  |  |  | 10 |  | 1 | 5.9e-07
100 V, 30 KHz |  | 99.9979 | 240 |  |  |  | 10 |  | 0.44097 | 6.8e-07
100 V, 50 KHz |  | 99.9966 | 240 |  |  |  | 20 |  | 0.83336 | 8.2e-07
100 V, 100 KHz |  | 99.991 | 600 |  |  |  | 30 |  | 2.0278 | 5.5e-07
100 V, 200 KHz |  | 99.9868 | 2600 |  |  |  | 50 |  | 1.8785 | 2.59e-06
1000 V, 55 Hz |  | 1000.018 | 79 |  |  |  | 15 |  | 1.2693 | 2.13e-06
1000 V, 300 Hz |  | 1000.025 | 79 |  |  |  | 15 |  | 2.4494 | 1.6e-06
700 V, 1 KHz |  | 699.997 | 80.714 |  |  |  | 50 |  | 1.1905 | 7.7e-07
DCI PERFORMANCE TEST
DCI Test | Expected Value | Measured A | Source Unc. ppm | Lower Limit A | Upper Limit A | Measured ppm | DUT Transfer STB ppm | Test Result % of SPEC | MEAS SDEV ppm
1 µADC | 1e-06 | 9.989e-07 | 10056.0617 | [5 empty] | 50.0551 | 1.224e-05
10 µADC | 1e-05 | 9.9993e-06 | 1045.0732 | [5 empty] | 0 | 8.2e-07
50 µADC | 5e-05 | 5.0001e-05 | 245 | [5 empty] | 0.99998 | 2.6e-07
100 µADC | 0.0001 | 0.0001 | 145 |  |  |  | 7 |  | 0 | 1.5e-07
-100 µADC | -0.0001 | -0.0001 | 145 |  |  |  | 7 |  | 0 | 1.2e-07
-50 µADC | -5e-05 | -5.0003e-05 | 245 | [5 empty] | 0 | 2.6e-07
-10 µADC | -1e-05 | -1.0001e-05 | -1044.8537 | [5 empty] | 0 | 1.16e-06
0.1 mADC | 0.0001 | 9.9996e-05 | 145 | [5 empty] | 5.0002 | 7.8e-07
0.5 mADC | 0.0005 | 0.00050001 | 65 | [5 empty] | 0.99998 | 2.3e-07
1 mADC | 0.001 | 0.001 | 55 |  |  |  | 7 |  | 0.49999 | 1e-07
-1 mADC | -0.001 | -0.001 | 55 |  |  |  | 7 |  | 0 | 1.4e-07
-0.5 mADC | -0.0005 | -0.00050002 | 65 | [5 empty] | -0.99997 | 2e-07
1 mADC | 0.001 | 0.00099928 | 55 | [5 empty] | 0 | 1.16e-06
5 mADC | 0.005 | 0.0049993 | 65 | [5 empty] | 0 | 2.8e-07
10 mADC | 0.01 | 0.0099993 | 55 |  |  |  | 7 |  | 0 | 1.2e-07
-10 mADC | -0.01 | -0.010001 | 55 |  |  |  | 7 |  | -0.49997 | 1.4e-07
-5 mADC | -0.005 | -0.0050007 | 65 | [5 empty] | -1.1545 | 2e-07
10 mADC | 0.01 | 0.0099984 | 55 | [5 empty] | 5.0008 | 8.5e-07
50 mADC | 0.05 | 0.049999 | 75 | [5 empty] | 1 | 2.4e-07
100 mADC | 0.1 | 0.099999 | 65 |  |  |  | 7 |  | 0 | 1.2e-07
-100 mADC | -0.1 | -0.1 | 65 |  |  |  | 7 |  | -0.57734 | 1.7e-07
-50 mADC | -0.05 | -0.050001 | 75 | [5 empty] | -0.99998 | 2.5e-07
0.1 ADC | 0.1 | 0.099998 | 65 | [5 empty] | 0 | 1.11e-06
0.5 ADC | 0.5 | 0.49999 | 135 | [5 empty] | 0 | 2.5e-07
-0.5 ADC | -0.5 | -0.50001 | 135 | [5 empty] | -0.99999 | 3e-07
-1.0 ADC | -1 | -0.99999 | 105 |  |  |  | 20 |  | 0 | 1e-07
1.0 ADC | 1 | 0.99997 | 105 |  |  |  | 20 |  | 0 | 1.2e-07
ACI PERFORMANCE TEST
ACI Test | Expected Value | Measured Aac | Source Unc. ppm | Lower Limit Aac | Upper Limit Aac | Measured ppm | DUT Transfer STB ppm | Test Result % of SPEC | MEAS SDEV ppm
10 µA, 10 Hz | 1e-05 | 9.9961e-06 | 3651.4241 | [5 empty] | 25.83 | 6.73e-06
10 µA, 20 Hz | 1e-05 | 9.9962e-06 | 2851.0977 | [5 empty] | 9.578 | 2.86e-06
10 µA, 30 Hz | 1e-05 | 9.9963e-06 | 2851.0549 | [5 empty] | 0 | 4.97e-06
10 µA, 40 Hz | 1e-05 | 9.9964e-06 | 2120.7635 | [5 empty] | 5.0018 | 3.76e-06
10 µA 55 Hz | 1e-05 | 9.9964e-06 | 2120.7635 | [5 empty] | 5.0018 | 5.58e-06
10 µA, 300 Hz | 1e-05 | 9.9967e-06 | 2120.7104 |  |  | -40 |  |  | 5.7754 | 5.93e-06
10 µA, 1 KHz | 1e-05 | 9.9967e-06 | 2120.7104 | [5 empty] | 5.7754 | 2.06e-05
10 µA, 5 KHz | 1e-05 | 9.9959e-06 | 5502.2559 | [5 empty] | 10.0041 | 6.69e-06
10 µA, 10 KHz | 1e-05 | 9.9942e-06 | 11506.6739 | [5 empty] | 5.0029 | 7.48e-06
100 µA, 10 Hz | 0.0001 | 0.00010002 | 950 | 9.9939e-05 | 0.00010006 | -40 | 50 |  | 5.6778 | 7.86e-06
100 µA, 20 Hz | 0.0001 | 0.00010001 | 600 |  |  |  | 50 |  | 2.0614 | 2.76e-06
100 µA, 30 Hz | 0.0001 | 0.00010001 | 600 | 9.9939e-05 | 0.00010006 |  | 50 |  | 1.4999 | 2.02e-06
100 µA, 40 Hz | 0.0001 | 0.00010001 | 320 |  |  |  | 50 |  | 1.4999 | 1.29e-06
100 µA 55 Hz | 0.0001 | 0.00010001 | 320 | 9.9939e-05 | 0.00010006 | -40 | 50 |  | 1.7077 | 2.36e-06
100 µA, 300 Hz | 0.0001 | 0.0001 | 320 |  |  |  | 50 |  | 2.5819 | 1.66e-06
100 µA, 1 KHz | 0.0001 | 0.0001 | 320 | 9.9939e-05 | 0.00010006 |  | 50 |  | 1.1547 | 2.39e-06
100 µA, 5 KHz | 0.0001 | 0.0001 | 0.1 |  |  |  | 100 |  | 0.49998 | 8.7e-07
100 µA, 10 KHz | 0.0001 | 0.00010002 | 0.25 | 9.9939e-05 | 0.00010006 |  | 300 |  | 0.57721 | 5.8e-07
1mA, 10 Hz | 0.001 | 0.0010002 | 700 |  |  |  | 40 |  | 4.9657 | 9.86e-06
1 mA, 20 Hz | 0.001 | 0.0010001 | 390 | 9.9939e-05 | 0.00010006 |  | 40 |  | 0.95737 | 1.63e-06
1 mA, 30 Hz | 0.001 | 0.0010001 | 390 |  |  |  | 40 |  | 0.99995 | 1.4e-06
1 mA, 40 Hz | 0.001 | 0.001 | 160 | 9.9939e-05 | 0.00010006 |  | 40 |  | 0.49998 | 1.53e-06
1 mA, 55 Hz | 0.001 | 0.001 | 160 |  |  |  | 40 |  | 3.0956 | 2.81e-06
1 mA, 300 Hz | 0.001 | 0.001 | 160 | 9.9939e-05 | 0.00010006 |  | 40 |  | 2.7537 | 1.8e-06
1 mA, 1 KHz | 0.001 | 0.001 | 160 |  |  |  | 40 |  | 1.2583 | 1.43e-06
1 mA, 5 KHz | 0.001 | 0.001 | 1000 | 9.9939e-05 | 0.00010006 |  | 70 |  | 0.57734 | 8.6e-07
1 mA, 10 KHz | 0.001 | 0.00099998 | 0.25 |  |  |  | 200 |  | 0.50001 | 6.3e-07
10 mA, 10 Hz | 0.01 | 0.010002 | 700 | 9.9939e-05 | 0.00010006 |  | 40 |  | 3.6962 | 9.04e-06
10 mA, 20 Hz | 0.01 | 0.010001 | 390 |  |  |  | 40 |  | 1.4999 | 2.02e-06
10 mA, 30 Hz | 0.01 | 0.010001 | 390 | 9.9939e-05 | 0.00010006 |  | 40 |  | 0.57732 | 1.15e-06
10 mA, 40 Hz | 0.01 | 0.01 | 160 |  |  |  | 40 |  | 0.99995 | 1.23e-06
10 mA, 55 Hz | 0.01 | 0.01 | 160 | 9.9939e-05 | 0.00010006 |  | 40 |  | 0.95739 | 2.59e-06
10 mA, 300 Hz | 0.01 | 0.01 | 160 |  |  |  | 40 |  | 1.4142 | 2.23e-06
10 mA, 1 KHz | 0.01 | 0.010001 | 160 | 9.9939e-05 | 0.00010006 |  | 40 |  | 1.2582 | 1.38e-06
10 mA, 5 KHz | 0.01 | 0.010001 | 1000 |  |  |  | 70 |  | 0.81638 | 6.9e-07
10 mA, 10 KHz | 0.01 | 0.0099999 | 0.25 | 9.9939e-05 | 0.00010006 |  | 200 |  | 0.50001 | 4e-07
100 mA, 10 Hz | 0.1 | 0.10002 | 700 |  |  |  | 40 |  | 6.1293 | 8.97e-06
100 mA, 20 Hz | 0.1 | 0.10001 | 390 | 9.9939e-05 | 0.00010006 |  | 40 |  | 1.7319 | 2.23e-06
100 mA, 30 Hz | 0.1 | 0.10001 | 390 |  |  |  | 40 |  | 0.95733 | 1.5e-06
100 mA, 40 Hz | 0.1 | 0.10001 | 160 | 9.9939e-05 | 0.00010006 |  | 40 |  | 1.2582 | 1.03e-06
100 mA, 55 Hz | 0.1 | 0.10001 | 160 |  |  |  | 40 |  | 2.6298 | 2.3e-06
100 mA, 300 Hz | 0.1 | 0.10001 | 160 | 9.9939e-05 | 0.00010006 |  | 40 |  | 2.0614 | 2.11e-06
100 mA, 1 KHz | 0.1 | 0.10001 | 160 |  |  |  | 40 |  | 2.0614 | 1.38e-06
100 mA, 5 KHz | 0.1 | 0.099997 | 0.1 | 9.9939e-05 | 0.00010006 |  | 70 |  | 0.57737 | 5.5e-07
100 mA, 10 KHz | 0.1 | 0.099989 | 0.25 |  |  |  | 200 |  | 0.95753 | 6e-07
1 A, 10 Hz | 1 | 0.99997 | -1 | 9.9939e-05 | 0.00010006 |  | 40 |  | 9.2874 | 1.028e-05
1 A, 20 Hz | 1 | 0.99997 | -1 |  |  |  | 40 |  | 2.0817 | 2.4e-06
1 A, 30 Hz | 1 | 0.99997 | -1 | 9.9939e-05 | 0.00010006 |  | 40 |  | 1.2583 | 1.17e-06
1 A, 40 Hz | 1 | 1.0001 | 640 |  |  |  | 40 |  | 1.7319 | 2.13e-06
1 A, 55 Hz | 1 | 1.0001 | 640 | 9.9939e-05 | 0.00010006 |  | 40 |  | 1.9147 | 2.98e-06
1 A, 300 Hz | 1 | 1.0001 | 640 |  |  |  | 40 |  | 0.81644 | 2.5e-06
1 A, 1 KHz | 1 | 1 | 640 | 9.9939e-05 | 0.00010006 |  | 40 |  | 0.81647 | 7.9e-07
1 A, 5 KHz | 1 | 0.99942 | 800 |  |  |  | 70 |  | 0 | 9.8e-07
1 A, 10 KHz | 1 | 0.99929 | 0.82 | 9.9939e-05 | 0.00010006 |  | 200 |  | 0.9581 | 8.5e-07
Test Summary | Initialization | DCV | OHM | ACV | DCI | ACI |  | Total
Time escape minute | 0.19 | 206.9 | 249.15 | 645.75 | 0 | 185.93