## Sheet
PERFORMANCE TEST
DUT
Manufacturer |  | Datron Instruments |  |  | S/N |  | 31255
Model Number |  | 4950 |  |  | Test date |  | 24-06-01 14:05:37
Meter Info |  | Datron 4950 STD |  |  | Ambient Temp. |  | 23.0
Options |  | 1,1,1 |  |  | CERT_AMB? |  | 23
Firmware |  | 890197-02.07 |  |  | TEMP? |  | 44.12
Date of Baseline Cal |  | "1999-SEP-16" |  |  | TEMP? BASE |  | 4.12
Date of Certified Cal |  | "2024-JUN-01" |  |  | TEMP? CERTIFIED |  | -0.02
CNFTST? |  | 0 |  |  | STANDBY? |  | "00:00"
DUT Confidence level |  | 30 Days Transfer Satbility
Calibrator
Reference standard |  | Options |  | Unc | Calibration Date |  | Due Date
Fluke 5700A |  |  |  | -- | -- |  | --
ADRmu12 |  | 10.000000 VDC |  | -- | -- |  | --
Fluke 742A-1 8927002 |  | 1.0000195 ohm |  | -- | -- |  | --
Fluke 742A-10k 7815006 |  | 10.000133 kohm
MFC
MFC Calibrate Date |  | 290524 |  |  | MFC last calibrated |  | 3 days
Calibrate Date WBFLAT |  | 100589 |  |  | MFC since WBFLAT |  | 0 days
Calibrate Date WB Gain |  | 250190 |  |  | MFC since WBGAIN |  | 0 days
MFC Calibrate Date Zero |  | 290524 |  |  | MFC since DCV Zero |  | 3 days
MFC Confidence level |  | 5700A 99% 90 Days |  |  | CAL CONST 11V Range +zero |  | 398.1793
CAL CONST 6.5V |  | +6.529890647135340E+00 |  |  | CAL CONST 11V Range -zero |  | 398.1784
CAL CONST 13V reference voltage |  | +1.306130764516422E+01 |  |  | CAL CONST, Zero calibration temperature |  | 23
CAL CONST 10KOHM true output resistance |  | +1.000045271113747E+04 |  |  | CAL CONST, ALL calibration temp |  | 23
CAL SONST 10kOHM standard resistance |  | +1.000069664541200E+04 |  |  | CAL_CONST? CHECK, WB_FLAT_TEMP |  | 23.9
CAL CONST 1 OHM |  | +9.996126071898848E-01 |  |  | CAL CONST,WBAND_TEMP
ETIME? |  | 6058 days 3.90 hours |  |  | CAL TEMP? CHECK
ONTIME? |  | 3 days 18.45 hours |  |  | CAL CONST,DACLIN |  | 0.359
DCV PERFORMANCE TEST
DCV Test VDC | Expected Value | Measured μV | Source Unc. μV | Lower Limit μV | Upper Limit μV | Deviation | DUT Transfer STB ppm | Test Result % of SPEC | MEAS SDEV μV
0.0mVDC | 0 | -0.02 | 7.5e-07 |  |  | N/A |  |  | 0.0092796 | 0.00626
0.0VDC | 0 | 4.7 | 7.5e-07 |  |  | N/A |  |  | 0.052705 | 0.01371
0.0VDC | 0 | 13 | 7.5e-07 |  |  | N/A |  |  | 0.52705 | 0.0672
0.0VDC | 0 | 260 | 7.5e-07 |  |  | N/A |  |  | 5 | 0.813
0.0VDC | 0 | -100 | 7.5e-07 |  |  | N/A |  |  | 0 | 4.34
DCV Test VDC | DMM Range | Measured V | Source Unc. ppm | Lower Limit V | Upper Limit V | Measured ppm | DUT Transfer STB ppm | Test Result % of SPEC | MEAS SDEV ppm
0.01 | 100mV | 0.0099999 | 81.5 | [5 empty] | 0.72649 | 6.7e-07
0.02 | 100mV | 0.02 | 44 | [5 empty] | 0.43302 | 2.8e-07
0.05 | 100mV | 0.05 | 21.5 | [5 empty] | 0.24495 | 1.1e-07
0.1 | 100mV | 0.1 | 14 |  |  |  | 3 |  | 0.072649 | 6e-08
-0.01 | 100mV | -0.01 | -81.5 | [5 empty] | -1.3333 | 5.6e-07
-0.02 | 100mV | -0.02 | -44 | [5 empty] | -0.35355 | 2.9e-07
-0.05 | 100mV | -0.05 | 21.5 | [5 empty] | -0.30185 | 1e-07
-0.1 | 100mV | -0.1 | 14 |  |  |  | 3 |  | -0.044096 | 6e-08
0.1 | 1V | 0.1 | 14 | [5 empty] | 1.1303 | 1.3e-07
0.2 | 1V | 0.2 | 10.25 | [5 empty] | 0.25 | 7e-08
0.5 | 1V | 0.5 | 5.9 | [5 empty] | 0.1 | 3e-08
1 | 1V | 1 | 4.7 |  |  |  | 1.5 |  | 0.1236 | 2e-08
-0.1 | 1V | -0.099995 | 14 | [5 empty] | -1.0541 | 1.4e-07
-0.2 | 1V | -0.19999 | 10.25 | [5 empty] | -0.16667 | 7e-08
-0.5 | 1V | -0.5 | 5.9 | [5 empty] | -0.088192 | 3e-08
-1 | 1V | -1 | 4.7 |  |  |  | 1.5 |  | -0.078174 | 2e-08
1 | 10V | 1 | 4.7 | [5 empty] | 0.5 | 8e-08
2 | 10V | 2 | 4.1 | [5 empty] | 0 | 4e-08
5 | 10V | 5 | 4.1 | [5 empty] | 0.10541 | 2e-08
10 | 10V | 10 | 3.8 |  |  |  | 1.5 |  | 0.060093 | 2e-08
19 | 10V | 18.9999 | 3.8158 |  |  |  | 1.5 |  | 0.057522 | 2e-08
-1 | 10V | -0.99999 | 4.7 | [5 empty] | -1.5366 | 7e-08
-2 | 10V | -2 | 4.1 | [5 empty] | -0.25 | 3e-08
-5 | 10V | -5 | 4.1 | [5 empty] | -0.066667 | 2e-08
-10 | 10V | -10 | 3.8 |  |  |  | 0.5 |  | -0.083333 | 2e-08
-19 | 10V | -19 | 3.8158 |  |  |  | 1.5 |  | -0.04884 | 1e-08
10 | 100V | 10.0002 | 3.8 | [5 empty] | 1.2247 | 7e-08
20 | 100V | 20.0002 | 3.8 | [5 empty] | 0.26352 | 6e-08
50 | 100V | 50.0001 | 7 | [5 empty] | 0.13333 | 3e-08
100 | 100V | 100 | 6 |  |  |  | 2 |  | 0.044096 | 2e-08
-10 | 100V | -9.9997 | 3.8 | [5 empty] | -0.7265 | 9e-08
-20 | 100V | -19.9997 | 3.8 | [5 empty] | -0.16667 | 4e-08
-50 | 100V | -49.9999 | 7 | [5 empty] | -0.13333 | 3e-08
-100 | 100V | -100 | 6 |  |  |  | 2 |  | -0.072648 | 2e-08
-100 | 1000V | -99.9999 | 6 | [5 empty] | -0.44096 | 5e-08
-200 | 1000V | -199.9995 | 5.5 | [5 empty] | 0 | 3e-08
-500 | 1000V | -500 | 8.2 | [5 empty] | -0.066667 | 2e-08
-1000 | 1000V | -1000.0001 | 7.6 |  |  |  | 2 |  | -0.033333 | 2e-08
100 | 1000V | 99.9998 | 6 | [5 empty] | 0.83333 | 6e-08
200 | 1000V | 199.9997 | 5.5 | [5 empty] | 0.16667 | 3e-08
500 | 1000V | 500.0002 | 8.2 | [5 empty] | 0.10541 | 2e-08
1000 | 1000V | 1000.0005 | 7.6 |  |  |  | 2 |  | 0.044096 | 1e-08
OHM PERFORMANCE TEST
OHM Test 4W | Calibrator ohm | Measured ohm | Source Unc. ppm | Lower Limit ohm | Upper Limit ohm | Measured ppm | DUT Transfer STB ppm | Test Result % of SPEC | MEAS SDEV ppm
1 Ω | 0.99961 | 0.9996 | 85 |  |  |  | 20 |  | 1.8567 | 1.69e-06
1.9 Ω | 1.9 | 1.8999 | 85 |  |  |  | 15 |  | 0.7797 | 7.8e-07
10 Ω | 10.0004 | 10.0003 | 26 |  |  |  | 5 |  | 0.10541 | 1.6e-07
19 Ω | 19.0039 | 19.0039 | 24 |  |  |  | 5 |  | 0.09766 | 8e-08
100 Ω | 100.0053 | 100.0056 | 15 |  |  |  | 3 |  | 0.052702 | 3e-08
190 Ω | 190.0032 | 190.0036 | 15 |  |  |  | 3 |  | 0.023208 | 2e-08
1 kΩ | 999.9469 | 999.9483 | 11 |  |  |  | 3 |  | 0.033335 | 3e-08
1.9 kΩ | 1899.9106 | 1899.9084 | 11 |  |  |  | 3 |  | 0.026317 | 2e-08
10 kΩ | 10000.453 | 10000.448 | 9 |  |  |  | 3 |  | 0.044094 | 3e-08
19 kΩ | 18999.722 | 18999.662 | 9 |  |  |  | 3 |  | 0.037217 | 2e-08
100 kΩ | 99998.75 | 99998.28 | 11 |  |  |  | 5 |  | 0.060094 | 2e-08
190 kΩ | 189994.66 | 189993.69 | 11 |  |  |  | 5 |  | 0.063257 | 1e-08
1 MΩ | 999908.7 | 999901.9 | 16 |  |  |  | 8 |  | 0.20618 | 3e-08
1.9 MΩ | 1899855.1 | 1899829.1 | 17 |  |  |  | 8 |  | 0.027742 | 2e-08
10MΩ | 9998564 | 9998405 | 33 |  |  |  | 12 |  | 0.11182 | 1e-07
19 MΩ | 18999500 | 18999897 | 43 |  |  |  | 12 |  | 0.28152 | 7e-08
100MΩ(2W) | 100009390 | 100011690 | 110 |  |  |  | 180 |  | 2.3439 | 1.7e-07
OHM Zero 4W | DUT Range | Measured ohm | Source Unc. ohm | Lower Limit ohm | Upper Limit ohm | Measured | DUT Transfer STB ppm | Test Result % of SPEC | MEAS SDEV ohm
0 Ω | 10 Ω | -8.2e-05 | 5e-05 |  |  | N/A |  |  | 1.0952e-05 | 1.531e-06
0 Ω | 100 Ω | 0.00054 | 5e-05 |  |  | N/A |  |  | 5.2705e-06 | 2.42e-06
0 Ω | 1 KΩ | 0.0056 | 5e-05 |  |  | N/A |  |  | 0 | 2.32e-05
0 Ω | 10 KΩ | 0.056 | 5e-05 |  |  | N/A |  |  | 0 | 0.000259
0 Ω | 100 KΩ | 0.13 | 5e-05 |  |  | N/A |  |  | 0 | 0.000695
0 Ω | 1 MΩ | 1.3 | 5e-05 |  |  | N/A |  |  | 0.033333 | 0.00524
0 Ω | 10 MΩ | 12 | 5e-05 |  |  | N/A |  |  | 0 | 0.0416
0 Ω | 100 MΩ(2W) | 130 | 5e-05 |  |  | N/A |  |  | 0 | 0.527
OHM Zero 2W | DUT Range | Measured ohm | Source Unc. ohm | Lower Limit ohm | Upper Limit ohm | Measured | DUT Transfer STB ppm | Test Result % of SPEC | MEAS SDEV ohm
0 Ω | 10 Ω | 0.41405 | 5e-05 |  |  | N/A |  |  | 0.00012071 | 1.976e-06
0 Ω | 100 Ω | 0.41076 | 5e-05 |  |  | N/A |  |  | 3.0046e-05 | 2.48e-06
0 Ω | 1 KΩ | 0.4158 | 5e-05 |  |  | N/A |  |  | 5e-05 | 2.34e-05
0 Ω | 10 KΩ | 0.467 | 5e-05 |  |  | N/A |  |  | 0.0005 | 0.000262
0 Ω | 100 KΩ | 0.54 | 5e-05 |  |  | N/A |  |  | 0 | 0.000675
0 Ω | 1 MΩ | 1.7 | 5e-05 |  |  | N/A |  |  | 0 | 0.00554
0 Ω | 10 MΩ | 12 | 5e-05 |  |  | N/A |  |  | 0.44096 | 0.0487
0 Ω | 100 MΩ | 130 | 5e-05 |  |  | N/A |  |  | 0 | 0.611
ACV PERFORMANCE TEST
ACV Linearity Checks | DUT Range | Measured Vac | Source Unc. ppm | Lower Limit Vac | Upper Limit Vac | Measured ppm of Input | DUT Transfer STB ppm | Test Result % of SPEC | MEAS SDEV ppm
1VAC/1kHz | 10 VAC | 0.99915 | 77 |  |  |  | 10 |  | 36.4315 | 3.51e-06
2VAC/1kHz | 10 VAC | 1.9995 | 73.5 |  |  |  | 10 |  | 6.6683 | 1.99e-06
5VAC/1kHz | 10 VAC | 4.9998 | 84 |  |  |  | 10 |  | 1.6667 | 1.46e-06
10VAC/1kHz | 10 VAC | 10 | 77 |  |  |  | 10 |  | 0.86603 | 1.09e-06
12VAC/1kHz | 10 VAC | 12 | 75.833 |  |  |  | 10 |  | 1.718 | 9.4e-07
15VAC/1kHz | 10 VAC | 15.0001 | 74.667 |  |  |  | 10 |  | 1.1276 | 7.9e-07
19VAC/1kHz | 10 VAC | 19.0002 | 73.684 |  |  |  | 10 |  | 0.9116 | 1.21e-06
ACV Performance Tset |  | Measured Vac | Source Unc. ppm | Lower Limit Vac | Upper Limit Vac | Measured ppm | DUT Transfer STB ppm | Test Result % of SPEC | MEAS SDEV ppm
0.001 V, 10 Hz |  | 0.00099568 | 5500 |  |  |  | 20 |  | 65.282 | 4.37e-06
0.001 V, 20 Hz |  | 0.00099578 | 5200 |  |  |  | 20 |  | 18.1814 | 5.62e-06
0.001 V, 30 Hz |  | 0.00099587 | 5200 |  |  |  | 20 |  | 10.0415 | 3.24e-06
0.001 V, 40 Hz |  | 0.00099585 | 5100 |  |  |  | 20 |  | 8.8559 | 5.32e-06
0.001 V, 55 Hz |  | 0.00099618 | 5100 |  |  |  | 20 |  | 5.0192 | 2.71e-06
0.001 V, 300 Hz |  | 0.00099615 | 5100 |  |  |  | 20 |  | 12.2948 | 2.68e-05
0.001 V, 1 KHz |  | 0.00099622 | 5100 |  |  |  | 20 |  | 14.1958 | 7.53e-06
0.001 V, 10 KHz |  | 0.00099623 | 5100 |  |  |  | 20 |  | 12.4071 | 8.1e-06
0.001 V, 20 KHz |  | 0.00099673 | 5100 |  |  |  | 20 |  | 10.1712 | 5.14e-06
0.001 V, 30 KHz |  | 0.00099667 | 5340 |  |  |  | 20 |  | 7.2891 | 3.63e-06
0.001 V, 50 KHz |  | 0.00099705 | 5340 |  |  |  | 20 |  | 8.358 | 6.53e-06
0.001 V, 100 KHz |  | 0.00099985 | 8800 |  |  |  | 50 |  | 4.4102 | 5.15e-06
0.001 V, 200 KHz |  | 0.00099711 | 16100 |  |  |  | 100 |  | 15.4104 | 8.56e-06
0.001 V, 300 KHz |  | 0.00099443 | 16100 |  |  |  | 100 |  | 11.243 | 4.23e-06
0.001 V, 500 KHz |  | 0.0010024 | 31500 |  |  |  | 200 |  | 24.1508 | 5.83e-05
0.001 V, 1 MHz |  | 0.0010099 | 33000 |  |  |  | 300 |  | 7.1936 | 1.563e-05
0.01 V, 10 Hz |  | 0.0099979 | 1100 |  |  |  | 20 |  | 11.1429 | 1.03e-06
0.01 V, 20 Hz |  | 0.0099969 | 800 |  |  |  | 20 |  | 1.6588 | 1.28e-06
0.01 V, 30 Hz |  | 0.0099967 | 800 |  |  |  | 20 |  | 0.97215 | 1.92e-06
0.01 V, 40 Hz |  | 0.0099964 | 700 |  |  |  | 20 |  | 0.60114 | 2.56e-06
0.01 V, 55 Hz |  | 0.0099965 | 700 |  |  |  | 20 |  | 1.0142 | 2.92e-06
0.01 V, 300 Hz |  | 0.0099964 | 700 |  |  |  | 20 |  | 1.3647 | 4.05e-06
0.01 V, 1 KHz |  | 0.0099965 | 700 |  |  |  | 20 |  | 1.4535 | 1.3e-06
0.01 V, 10 KHz |  | 0.0099976 | 700 |  |  |  | 20 |  | 0.88213 | 1.21e-06
0.01 V, 20 KHz |  | 0.0099986 | 700 |  |  |  | 20 |  | 1.3644 | 8e-07
0.01 V, 30 KHz |  | 0.0099989 | 940 |  |  |  | 20 |  | 0.60099 | 5.2e-07
0.01 V, 50 KHz |  | 0.01 | 940 |  |  |  | 20 |  | 1.0138 | 7.1e-07
0.01 V, 100 KHz |  | 0.010005 | 1600 |  |  |  | 50 |  | 1.0924 | 8.6e-07
0.01 V, 200 KHz |  | 0.0099973 | 2600 |  |  |  | 100 |  | 1.3948 | 8.7e-07
0.01 V, 300 KHz |  | 0.0099945 | 2600 |  |  |  | 100 |  | 0.60126 | 8.4e-07
0.01 V, 500 KHz |  | 0.010006 | 4500 |  |  |  | 200 |  | 2.4874 | 5.87e-06
0.01 V, 1 MHz |  | 0.010056 | 6000 |  |  |  | 300 |  | 1.0482 | 1.86e-06
0.1 V, 10 Hz |  | 0.10002 | 660 |  |  |  | 20 |  | 8.3504 | 8.54e-06
0.1 V, 20 Hz |  | 0.1 | 300 |  |  |  | 20 |  | 1.5811 | 2.41e-06
0.1 V, 30 Hz |  | 0.1 | 300 |  |  |  | 20 |  | 1.3017 | 1.59e-06
0.1 V, 40 Hz |  | 0.1 | 195 |  |  |  | 20 |  | 1.2693 | 1.86e-06
0.1 V, 55 Hz |  | 0.1 | 195 |  |  |  | 20 |  | 1.2247 | 2.04e-06
0.1 V, 300 Hz |  | 0.1 | 195 |  |  |  | 20 |  | 1.7321 | 2.41e-06
0.1 V, 1 KHz |  | 0.1 | 195 |  |  |  | 20 |  | 0.83333 | 1.22e-06
0.1 V, 10 KHz |  | 0.1 | 195 |  |  |  | 20 |  | 0.88192 | 6.3e-07
0.1 V, 20 KHz |  | 0.1 | 195 |  |  |  | 20 |  | 0.70711 | 4.5e-07
0.1 V, 30 KHz |  | 0.1 | 400 |  |  |  | 20 |  | 0.92796 | 6.4e-07
0.1 V, 50 KHz |  | 0.1 | 400 |  |  |  | 20 |  | 0.78174 | 5.4e-07
0.1 V, 100 KHz |  | 0.1 | 1050 |  |  |  | 50 |  | 1.2693 | 5.9e-07
0.1 V, 200 KHz |  | 0.10002 | 1240 |  |  |  | 100 |  | 0.70697 | 7.2e-07
0.1 V, 300 KHz |  | 0.099993 | 1240 |  |  |  | 100 |  | 0.60097 | 5.6e-07
0.1 V, 500 KHz |  | 0.099957 | 1900 |  |  |  | 200 |  | 0.50021 | 8.7e-07
0.1 V, 1 MHz |  | 0.10005 | 4000 |  |  |  | 300 |  | 2.1462 | 1.08e-06
1 V, 10 Hz |  | 1.0001 | 600 |  |  |  | 10 |  | 8.6658 | 8.4e-06
1 V, 20 Hz |  | 1 | 180 |  |  |  | 10 |  | 1.3944 | 2.42e-06
1 V, 30 Hz |  | 1 | 180 |  |  |  | 10 |  | 1.1304 | 9.2e-07
1 V, 40 Hz |  | 0.99999 | 77 |  |  |  | 10 |  | 0.70712 | 1.78e-06
1 V, 55 Hz |  | 0.99999 | 77 |  |  |  | 10 |  | 1.5093 | 1.44e-06
1 V, 300 Hz |  | 0.99999 | 77 |  |  |  | 10 |  | 0.70711 | 1.92e-06
1 V, 1 KHz |  | 1 | 77 |  |  |  | 10 |  | 1.4142 | 8.9e-07
1 V, 10 KHz |  | 0.99999 | 77 |  |  |  | 10 |  | 0.72649 | 6e-07
1 V, 20 KHz |  | 1 | 77 |  |  |  | 10 |  | 0.44095 | 5e-07
1 V, 30 KHz |  | 1 | 140 |  |  |  | 10 |  | 1.0138 | 4.4e-07
1 V, 50 KHz |  | 1.0001 | 140 |  |  |  | 20 |  | 1.118 | 4.3e-07
1 V, 100 KHz |  | 0.9998 | 310 |  |  |  | 30 |  | 0.8662 | 6.7e-07
1 V, 200 KHz |  | 0.99936 | 550 |  |  |  | 70 |  | 0.70756 | 5.4e-07
1 V, 300 KHz |  | 0.99966 | 550 |  |  |  | 70 |  | 1.1308 | 4e-07
1 V, 500 KHz |  | 0.99897 | 1400 |  |  |  | 100 |  | 0.50051 | 4.3e-07
1 V, 1 MHz |  | 0.99953 | 3000 |  |  |  | 200 |  | 1.9012 | 1.42e-06
10 V, 10 Hz |  | 10.0011 | 600 |  |  |  | 10 |  | 13.2273 | 1.021e-05
10 V, 20 Hz |  | 10.0001 | 180 |  |  |  | 10 |  | 1.7873 | 2.34e-06
10 V, 30 Hz |  | 9.9999 | 180 |  |  |  | 10 |  | 1.3642 | 1.54e-06
10 V, 40 Hz |  | 9.9999 | 77 |  |  |  | 10 |  | 0.78175 | 2.47e-06
10 V, 55 Hz |  | 9.9998 | 77 |  |  |  | 10 |  | 0.78175 | 2.28e-06
10 V, 300 Hz |  | 9.9998 | 77 |  |  |  | 10 |  | 1.7321 | 1.72e-06
10 V, 1 KHz |  | 10 | 77 |  |  |  | 10 |  | 1.3333 | 1.34e-06
10 V, 10 KHz |  | 9.9998 | 77 |  |  |  | 10 |  | 0.88194 | 2.9e-07
10 V, 20 KHz |  | 10.0001 | 77 |  |  |  | 10 |  | 0.66666 | 8.6e-07
10 V, 30 KHz |  | 10.0001 | 140 |  |  |  | 10 |  | 0.86601 | 6.6e-07
10 V, 50 KHz |  | 10.0003 | 140 |  |  |  | 20 |  | 0.66665 | 6.2e-07
10 V, 100 KHz |  | 9.9981 | 270 |  |  |  | 30 |  | 1.014 | 9.2e-07
10 V, 200 KHz |  | 9.9932 | 670 |  |  |  | 70 |  | 0.44126 | 6.3e-07
10 V, 300 KHz |  | 9.9961 | 670 |  |  |  | 70 |  | 1.3022 | 7.5e-07
10 V, 500 KHz |  | 9.9892 | 1700 |  |  |  | 100 |  | 0.86696 | 3.5e-07
10 V, 1 MHz |  | 10.0027 | 3500 |  |  |  | 200 |  | 1.0926 | 2.36e-06
19 V, 10 Hz |  | 19.0066 | 552.63 |  |  |  | 10 |  | 5.9351 | 6.37e-06
19 V, 20 Hz |  | 19.0016 | 165.79 |  |  |  | 10 |  | 1.0191 | 1.62e-06
19 V, 30 Hz |  | 19.0008 | 165.79 |  |  |  | 10 |  | 0.38234 | 9.1e-07
19 V, 40 Hz |  | 19.0004 | 73.684 |  |  |  | 10 |  | 0.53357 | 3.27e-06
19 V, 55 Hz |  | 19.0001 | 73.684 |  |  |  | 10 |  | 0.76472 | 2.3e-06
19 V, 300 Hz |  | 18.9999 | 73.684 |  |  |  | 10 |  | 1.5593 | 1.75e-06
19 V, 1 KHz |  | 19 | 73.684 |  |  |  | 10 |  | 1.0886 | 1.12e-06
19 V, 10 KHz |  | 18.9999 | 73.684 |  |  |  | 10 |  | 0.41144 | 4.8e-07
19 V, 20 KHz |  | 19.0004 | 73.684 |  |  |  | 10 |  | 1.0563 | 5.3e-07
19 V, 30 KHz |  | 19.0006 | 130.53 |  |  |  | 10 |  | 0.64458 | 4.6e-07
19 V, 50 KHz |  | 19.0009 | 130.53 |  |  |  | 20 |  | 0.45578 | 2.2e-07
19 V, 100 KHz |  | 18.9969 | 251.05 |  |  |  | 30 |  | 0.59504 | 3.6e-07
19 V, 200 KHz |  | 18.9879 | 589.47 |  |  |  | 70 |  | 1.2099 | 2.9e-07
19 V, 300 KHz |  | 18.9937 | 589.47 |  |  |  | 70 |  | 1.6672 | 4.9e-07
19 V, 500 KHz |  | 18.9812 | 1463.2 |  |  |  | 100 |  | 1.8059 | 3.3e-07
19 V, 1 MHz |  | 19.0105 | 3073.7 |  |  |  | 200 |  | 2.6879 | 1.5e-06
100 V, 10 Hz |  | 100.0108 | 600 |  |  |  | 10 |  | 6.3129 | 6.79e-06
100 V, 20 Hz |  | 100.0009 | 180 |  |  |  | 10 |  | 0.92795 | 2.61e-06
100 V, 30 Hz |  | 99.9995 | 180 |  |  |  | 10 |  | 0.92797 | 1.48e-06
100 V, 40 Hz |  | 99.9983 | 85 |  |  |  | 10 |  | 1.2248 | 2.29e-06
100 V, 55 Hz |  | 99.9984 | 85 |  |  |  | 10 |  | 0.72649 | 1.94e-06
100 V, 300 Hz |  | 99.9978 | 85 |  |  |  | 10 |  | 0.60094 | 2.12e-06
100 V, 1 KHz |  | 99.9973 | 85 |  |  |  | 10 |  | 0.88194 | 1.08e-06
100 V, 10 KHz |  | 99.997 | 85 |  |  |  | 10 |  | 0.78176 | 3.2e-07
100 V, 20 KHz |  | 99.997 | 85 |  |  |  | 10 |  | 0.78176 | 8.2e-07
100 V, 30 KHz |  | 99.9978 | 240 |  |  |  | 10 |  | 1.3017 | 2.9e-07
100 V, 50 KHz |  | 99.997 | 240 |  |  |  | 20 |  | 0.97185 | 6.7e-07
100 V, 100 KHz |  | 99.9927 | 600 |  |  |  | 30 |  | 1.6416 | 9.3e-07
100 V, 200 KHz |  | 99.9948 | 2600 |  |  |  | 50 |  | 2.5713 | 4.29e-06
1000 V, 55 Hz |  | 1000.023 | 79 |  |  |  | 15 |  | 0.86601 | 2.13e-06
1000 V, 300 Hz |  | 1000.032 | 79 |  |  |  | 15 |  | 1.5092 | 1.83e-06
700 V, 1 KHz |  | 700.006 | 80.714 |  |  |  | 50 |  | 0.71428 | 1.15e-06
DCI PERFORMANCE TEST
DCI Test | Expected Value | Measured A | Source Unc. ppm | Lower Limit A | Upper Limit A | Measured ppm | DUT Transfer STB ppm | Test Result % of SPEC | MEAS SDEV ppm
1 µADC | 1e-06 | 9.99e-07 | 10055.0551 | [5 empty] | 0 | 1.204e-05
10 µADC | 1e-05 | 9.9995e-06 | 1045.0523 | [5 empty] | 0 | 1.07e-06
50 µADC | 5e-05 | 5.0002e-05 | 245 | [5 empty] | 0 | 2e-07
100 µADC | 0.0001 | 0.0001 | 145 |  |  |  | 7 |  | 0.49998 | 1.1e-07
-100 µADC | -0.0001 | -0.00010001 | 145 |  |  |  | 7 |  | 0 | 1.1e-07
-50 µADC | -5e-05 | -5.0004e-05 | 245 | [5 empty] | 0 | 2.2e-07
-10 µADC | -1e-05 | -1.0002e-05 | -1044.8328 | [5 empty] | 0 | 8.8e-07
0.1 mADC | 0.0001 | 9.9998e-05 | 145 | [5 empty] | 0 | 1.2e-06
0.5 mADC | 0.0005 | 0.00050001 | 65 | [5 empty] | 0 | 1.8e-07
1 mADC | 0.001 | 0.001 | 55 |  |  |  | 7 |  | 0 | 1.3e-07
-1 mADC | -0.001 | -0.001 | 55 |  |  |  | 7 |  | -0.57733 | 1e-07
-0.5 mADC | -0.0005 | -0.00050003 | 65 | [5 empty] | -1.1546 | 2e-07
1 mADC | 0.001 | 0.001 | 55 | [5 empty] | 0 | 1.24e-06
5 mADC | 0.005 | 0.0050001 | 65 | [5 empty] | 0.99998 | 2.7e-07
10 mADC | 0.01 | 0.01 | 55 |  |  |  | 7 |  | 0 | 1.1e-07
-10 mADC | -0.01 | -0.01 | 55 |  |  |  | 7 |  | -0.5 | 1.1e-07
-5 mADC | -0.005 | -0.005 | 65 | [5 empty] | -1 | 2.4e-07
10 mADC | 0.01 | 0.0099987 | 55 | [5 empty] | 0 | 1.36e-06
50 mADC | 0.05 | 0.05 | 75 | [5 empty] | 1 | 2e-07
100 mADC | 0.1 | 0.1 | 65 |  |  |  | 7 |  | 0 | 1.2e-07
-100 mADC | -0.1 | -0.1 | 65 |  |  |  | 7 |  | -0.49998 | 1.7e-07
-50 mADC | -0.05 | -0.050002 | 75 | [5 empty] | 0 | 2.5e-07
0.1 ADC | 0.1 | 0.1 | 65 | [5 empty] | 0 | 1.26e-06
0.5 ADC | 0.5 | 0.50003 | 135 | [5 empty] | 0 | 2.1e-07
-0.5 ADC | -0.5 | -0.50004 | 135 | [5 empty] | -0.99993 | 1.9e-07
-1.0 ADC | -1 | -1 | 105 |  |  |  | 20 |  | 0 | 1.1e-07
1.0 ADC | 1 | 1 | 105 |  |  |  | 20 |  | 0.57733 | 9e-08
ACI PERFORMANCE TEST
ACI Test | Expected Value | Measured Aac | Source Unc. ppm | Lower Limit Aac | Upper Limit Aac | Measured ppm | DUT Transfer STB ppm | Test Result % of SPEC | MEAS SDEV ppm
10 µA, 10 Hz | 1e-05 | 9.9998e-06 | 3650.073 | [5 empty] | 8.1651 | 9.25e-06
10 µA, 20 Hz | 1e-05 | 9.9995e-06 | 2850.1425 | [5 empty] | 5.0003 | 4.57e-06
10 µA, 30 Hz | 1e-05 | 9.9992e-06 | 2850.2423 | [5 empty] | 5.774 | 4.11e-06
10 µA, 40 Hz | 1e-05 | 9.9993e-06 | 2120.1484 | [5 empty] | 5.0004 | 2e-06
10 µA 55 Hz | 1e-05 | 9.9992e-06 | 2120.1802 | [5 empty] | 5.774 | 3.82e-06
10 µA, 300 Hz | 1e-05 | 9.9992e-06 | 2120.1696 |  |  | -40 |  |  | 5.0004 | 6.07e-06
10 µA, 1 KHz | 1e-05 | 9.999e-06 | 2120.2226 | [5 empty] | 18.9317 | 2.65e-05
10 µA, 5 KHz | 1e-05 | 9.9974e-06 | 5501.4304 | [5 empty] | 17.325 | 9.82e-06
10 µA, 10 KHz | 1e-05 | 9.9958e-06 | 11504.832 | [5 empty] | 12.5883 | 7.28e-06
100 µA, 10 Hz | 0.0001 | 0.00010002 | 950 | 9.9939e-05 | 0.00010006 | -40 | 50 |  | 6.1901 | 8.8e-06
100 µA, 20 Hz | 0.0001 | 0.00010001 | 600 |  |  |  | 50 |  | 0.49995 | 2.79e-06
100 µA, 30 Hz | 0.0001 | 0.00010001 | 600 | 9.9939e-05 | 0.00010006 |  | 50 |  | 0.95734 | 1.84e-06
100 µA, 40 Hz | 0.0001 | 0.00010001 | 320 |  |  |  | 50 |  | 0.81643 | 2.57e-06
100 µA 55 Hz | 0.0001 | 0.00010001 | 320 | 9.9939e-05 | 0.00010006 | -40 | 50 |  | 1.4141 | 1.29e-06
100 µA, 300 Hz | 0.0001 | 0.00010001 | 320 |  |  |  | 50 |  | 1.4141 | 2.12e-06
100 µA, 1 KHz | 0.0001 | 0.0001 | 320 | 9.9939e-05 | 0.00010006 |  | 50 |  | 2.4999 | 2.61e-06
100 µA, 5 KHz | 0.0001 | 0.0001 | 0.1 |  |  |  | 100 |  | 0.49998 | 4.6e-07
100 µA, 10 KHz | 0.0001 | 0.00010002 | 0.25 | 9.9939e-05 | 0.00010006 |  | 300 |  | 0.95722 | 6e-07
1mA, 10 Hz | 0.001 | 0.0010002 | 700 |  |  |  | 40 |  | 6.6509 | 8.7e-06
1 mA, 20 Hz | 0.001 | 0.0010001 | 390 | 9.9939e-05 | 0.00010006 |  | 40 |  | 0.5773 | 2.44e-06
1 mA, 30 Hz | 0.001 | 0.0010001 | 390 |  |  |  | 40 |  | 0.95736 | 9.9e-07
1 mA, 40 Hz | 0.001 | 0.0010001 | 160 | 9.9939e-05 | 0.00010006 |  | 40 |  | 0.57732 | 1.74e-06
1 mA, 55 Hz | 0.001 | 0.0010001 | 160 |  |  |  | 40 |  | 0.95737 | 2.49e-06
1 mA, 300 Hz | 0.001 | 0.001 | 160 | 9.9939e-05 | 0.00010006 |  | 40 |  | 0.81646 | 1.98e-06
1 mA, 1 KHz | 0.001 | 0.001 | 160 |  |  |  | 40 |  | 0.49998 | 1.38e-06
1 mA, 5 KHz | 0.001 | 0.001 | 1000 | 9.9939e-05 | 0.00010006 |  | 70 |  | 0.57733 | 5.1e-07
1 mA, 10 KHz | 0.001 | 0.001 | 0.25 |  |  |  | 200 |  | 0.57735 | 2.9e-07
10 mA, 10 Hz | 0.01 | 0.010002 | 700 | 9.9939e-05 | 0.00010006 |  | 40 |  | 5.2905 | 1.076e-05
10 mA, 20 Hz | 0.01 | 0.010001 | 390 |  |  |  | 40 |  | 0.49996 | 3.01e-06
10 mA, 30 Hz | 0.01 | 0.010001 | 390 | 9.9939e-05 | 0.00010006 |  | 40 |  | 1.4141 | 1.11e-06
10 mA, 40 Hz | 0.01 | 0.010001 | 160 |  |  |  | 40 |  | 0.49997 | 1.96e-06
10 mA, 55 Hz | 0.01 | 0.010001 | 160 | 9.9939e-05 | 0.00010006 |  | 40 |  | 0.95737 | 1.73e-06
10 mA, 300 Hz | 0.01 | 0.010001 | 160 |  |  |  | 40 |  | 0.57732 | 2.25e-06
10 mA, 1 KHz | 0.01 | 0.010001 | 160 | 9.9939e-05 | 0.00010006 |  | 40 |  | 1.1546 | 1.91e-06
10 mA, 5 KHz | 0.01 | 0.010002 | 1000 |  |  |  | 70 |  | 0.49992 | 4.3e-07
10 mA, 10 KHz | 0.01 | 0.01 | 0.25 | 9.9939e-05 | 0.00010006 |  | 200 |  | 0.5 | 4.8e-07
100 mA, 10 Hz | 0.1 | 0.10002 | 700 |  |  |  | 40 |  | 4.6177 | 1.065e-05
100 mA, 20 Hz | 0.1 | 0.10001 | 390 | 9.9939e-05 | 0.00010006 |  | 40 |  | 0.81638 | 2.22e-06
100 mA, 30 Hz | 0.1 | 0.10001 | 390 |  |  |  | 40 |  | 0.57728 | 1.36e-06
100 mA, 40 Hz | 0.1 | 0.10001 | 160 | 9.9939e-05 | 0.00010006 |  | 40 |  | 0.49995 | 2.25e-06
100 mA, 55 Hz | 0.1 | 0.10001 | 160 |  |  |  | 40 |  | 1.2909 | 1.56e-06
100 mA, 300 Hz | 0.1 | 0.10001 | 160 | 9.9939e-05 | 0.00010006 |  | 40 |  | 0.81641 | 1.97e-06
100 mA, 1 KHz | 0.1 | 0.10001 | 160 |  |  |  | 40 |  | 1.7076 | 1.21e-06
100 mA, 5 KHz | 0.1 | 0.1 | 0.1 | 9.9939e-05 | 0.00010006 |  | 70 |  | 0.8165 | 3.7e-07
100 mA, 10 KHz | 0.1 | 0.099995 | 0.25 |  |  |  | 200 |  | 1.1548 | 3.2e-07
1 A, 10 Hz | 1 | 1 | -1 | 9.9939e-05 | 0.00010006 |  | 40 |  | 13.2659 | 9.55e-06
1 A, 20 Hz | 1 | 1 | -1 |  |  |  | 40 |  | 2.0816 | 2.07e-06
1 A, 30 Hz | 1 | 1 | -1 | 9.9939e-05 | 0.00010006 |  | 40 |  | 1.2909 | 1.29e-06
1 A, 40 Hz | 1 | 1.0001 | 640 |  |  |  | 40 |  | 0.99985 | 1.83e-06
1 A, 55 Hz | 1 | 1.0001 | 640 | 9.9939e-05 | 0.00010006 |  | 40 |  | 0.81638 | 1.69e-06
1 A, 300 Hz | 1 | 1.0001 | 640 |  |  |  | 40 |  | 0.81638 | 1.74e-06
1 A, 1 KHz | 1 | 1.0001 | 640 | 9.9939e-05 | 0.00010006 |  | 40 |  | 0.81641 | 8.8e-07
1 A, 5 KHz | 1 | 0.9994 | 800 |  |  |  | 70 |  | 1.5009 | 4.6e-07
1 A, 10 KHz | 1 | 0.99973 | 0.82 | 9.9939e-05 | 0.00010006 |  | 200 |  | 0.95769 | 4.3e-07
Test Summary | Initialization | DCV | OHM | ACV | DCI | ACI |  | Total
Time escape minute | 0.19 | 206.89 | 249.13 | 662.59 | 40.39 | 185.16