## Sheet
PERFORMANCE TEST
DUT
Manufacturer |  | Datron Instruments |  |  | S/N |  | 31255
Model Number |  | 4950 |  |  | Test date |  | 23-09-24 18:11:25
Meter Info |  | Datron 4950 STD |  |  | Ambient Temp. |  | 23.0
Options |  | 1,1,1 |  |  | CERT_AMB? |  | 23
Firmware |  | 890197-02.07 |  |  | TEMP? |  | 43.24
Date of Baseline Cal |  | "1999-SEP-16" |  |  | TEMP? BASE |  | 3.25
Date of Certified Cal |  | "2023-SEP-24" |  |  | TEMP? CERTIFIED |  | 0.06
CNFTST? |  | 0 |  |  | STANDBY? |  | "00:00"
DUT Confidence level |  | 30 Days Transfer Satbility
Calibrator
Reference standard |  | Options |  | Unc | Calibration Date |  | Due Date
Fluke 5700A |  |  |  | -- | -- |  | --
ADRmu12 |  | 10.000000 VDC |  | -- | -- |  | --
Fluke 742A-1 8927002 |  | 1.0000195 ohm |  | -- | -- |  | --
Fluke 742A-10k 7815006 |  | 10.000133 kohm
MFC
MFC Calibrate Date |  | 030622 |  |  | MFC last calibrated |  | 478 days
Calibrate Date WBFLAT |  | 100589 |  |  | MFC since WBFLAT |  | 0 days
Calibrate Date WB Gain |  | 250190 |  |  | MFC since WBGAIN |  | 0 days
MFC Calibrate Date Zero |  | 240923 |  |  | MFC since DCV Zero |  | 0 days
MFC Confidence level |  | 5700A 99% 90 Days |  |  | CAL CONST 11V Range +zero |  | 398.1796
CAL CONST 6.5V |  | +6.529921596265443E+00 |  |  | CAL CONST 11V Range -zero |  | 398.1789
CAL CONST 13V reference voltage |  | +1.306136884003510E+01 |  |  | CAL CONST, Zero calibration temperature |  | 23
CAL CONST 10KOHM true output resistance |  | +1.000039902274076E+04 |  |  | CAL CONST, ALL calibration temp |  | 23
CAL SONST 10kOHM standard resistance |  | +1.000059914344082E+04 |  |  | CAL_CONST? CHECK, WB_FLAT_TEMP |  | 23.9
CAL CONST 1 OHM |  | +9.996206897944441E-01 |  |  | CAL CONST,WBAND_TEMP
ETIME? |  | 6052 days 5.28 hours |  |  | CAL TEMP? CHECK
ONTIME? |  | 0 days 0.78 hours |  |  | CAL CONST,DACLIN |  | 0.359
DCV PERFORMANCE TEST
DCV Test VDC | Expected Value | Measured μV | Source Unc. μV | Lower Limit μV | Upper Limit μV | Deviation | DUT Transfer STB ppm | Test Result % of SPEC | MEAS SDEV μV
0.0mVDC | 0 | 0.03 | 7.5e-07 |  |  | N/A |  |  | 0.015092 | 0.00622
0.0VDC | 0 | 0.7 | 7.5e-07 |  |  | N/A |  |  | 0.044096 | 0.01317
0.0VDC | 0 | 4 | 7.5e-07 |  |  | N/A |  |  | 0 | 0.0754
0.0VDC | 0 | 10 | 7.5e-07 |  |  | N/A |  |  | 0 | 0.744
0.0VDC | 0 | 300 | 7.5e-07 |  |  | N/A |  |  | 0 | 4.36
DCV Test VDC | DMM Range | Measured V | Source Unc. ppm | Lower Limit V | Upper Limit V | Measured ppm | DUT Transfer STB ppm | Test Result % of SPEC | MEAS SDEV ppm
0.01 | 100mV | 0.01 | 81.5 | [5 empty] | 0.78173 | 6.6e-07
0.02 | 100mV | 0.02 | 44 | [5 empty] | 0.56519 | 3e-07
0.05 | 100mV | 0.05 | 21.5 | [5 empty] | 0.24495 | 1.3e-07
0.1 | 100mV | 0.1 | 14 |  |  |  | 3 |  | 0.14142 | 6e-08
-0.01 | 100mV | -0.0099999 | -81.5 | [5 empty] | -1.3229 | 6.3e-07
-0.02 | 100mV | -0.02 | -44 | [5 empty] | -0.60093 | 2.6e-07
-0.05 | 100mV | -0.05 | 21.5 | [5 empty] | -0.24037 | 1.2e-07
-0.1 | 100mV | -0.1 | 14 |  |  |  | 3 |  | -0.1 | 6e-08
0.1 | 1V | 0.1 | 14 | [5 empty] | 1.0929 | 1.3e-07
0.2 | 1V | 0.2 | 10.25 | [5 empty] | 0.22048 | 7e-08
0.5 | 1V | 0.5 | 5.9 | [5 empty] | 0.088192 | 3e-08
1 | 1V | 1 | 4.7 |  |  |  | 1.5 |  | 0.083333 | 2e-08
-0.1 | 1V | -0.099999 | 14 | [5 empty] | -1 | 1.2e-07
-0.2 | 1V | -0.2 | 10.25 | [5 empty] | -0.26352 | 6e-08
-0.5 | 1V | -0.5 | 5.9 | [5 empty] | -0.1 | 3e-08
-1 | 1V | -1 | 4.7 |  |  |  | 1.5 |  | -0.083333 | 2e-08
1 | 10V | 1 | 4.7 | [5 empty] | 0.72648 | 9e-08
2 | 10V | 2 | 4.1 | [5 empty] | 0 | 5e-08
5 | 10V | 5 | 4.1 | [5 empty] | 0.10541 | 2e-08
10 | 10V | 10 | 3.8 |  |  |  | 1.5 |  | 0.070711 | 2e-08
19 | 10V | 19 | 3.8158 |  |  |  | 1.5 |  | 0.057521 | 1e-08
-1 | 10V | -0.99999 | 4.7 | [5 empty] | -1.5366 | 7e-08
-2 | 10V | -2 | 4.1 | [5 empty] | -0.25 | 4e-08
-5 | 10V | -5 | 4.1 | [5 empty] | -0.066667 | 2e-08
-10 | 10V | -10 | 3.8 |  |  |  | 0.5 |  | -0.060093 | 2e-08
-19 | 10V | -19 | 3.8158 |  |  |  | 1.5 |  | -0.059494 | 2e-08
10 | 100V | 10 | 3.8 | [5 empty] | 1.6159 | 7e-08
20 | 100V | 20 | 3.8 | [5 empty] | 0.25 | 5e-08
50 | 100V | 50 | 7 | [5 empty] | 0.088192 | 3e-08
100 | 100V | 100 | 6 |  |  |  | 2 |  | 0.072648 | 1e-08
-10 | 100V | -10 | 3.8 | [5 empty] | -0.86603 | 9e-08
-20 | 100V | -20 | 3.8 | [5 empty] | 0 | 4e-08
-50 | 100V | -50 | 7 | [5 empty] | 0 | 3e-08
-100 | 100V | -100 | 6 |  |  |  | 2 |  | -0.072648 | 2e-08
-100 | 1000V | -99.9998 | 6 | [5 empty] | -0.33333 | 4e-08
-200 | 1000V | -199.9996 | 5.5 | [5 empty] | -0.22048 | 3e-08
-500 | 1000V | -500.0009 | 8.2 | [5 empty] | -0.1 | 2e-08
-1000 | 1000V | -1000.002 | 7.6 |  |  |  | 2 |  | -0.052705 | 1e-08
100 | 1000V | 100.0003 | 6 | [5 empty] | 0.7071 | 5e-08
200 | 1000V | 200.0004 | 5.5 | [5 empty] | 0.22048 | 3e-08
500 | 1000V | 500.0013 | 8.2 | [5 empty] | 0.088191 | 2e-08
1000 | 1000V | 1000.0024 | 7.6 |  |  |  | 2 |  | 0 | 1e-08
OHM PERFORMANCE TEST
OHM Test 4W | Calibrator ohm | Measured ohm | Source Unc. ppm | Lower Limit ohm | Upper Limit ohm | Measured ppm | DUT Transfer STB ppm | Test Result % of SPEC | MEAS SDEV ppm
1 Ω | 0.99962 | 0.99958 | 85 |  |  |  | 20 |  | 2.2983 | 1.82e-06
1.9 Ω | 1.9 | 1.8999 | 85 |  |  |  | 15 |  | 0.78952 | 9.5e-07
10 Ω | 10.0003 | 10.0003 | 26 |  |  |  | 5 |  | 0.13228 | 1.6e-07
19 Ω | 19.0037 | 19.0038 | 24 |  |  |  | 5 |  | 0.086376 | 9e-08
100 Ω | 100.0049 | 100.0052 | 15 |  |  |  | 3 |  | 0.052702 | 3e-08
190 Ω | 190.0021 | 190.0031 | 15 |  |  |  | 3 |  | 0 | 2e-08
1 kΩ | 999.9422 | 999.9451 | 11 |  |  |  | 3 |  | 0.050003 | 3e-08
1.9 kΩ | 1899.9031 | 1899.9048 | 11 |  |  |  | 3 |  | 0.017545 | 2e-08
10 kΩ | 10000.399 | 10000.415 | 9 |  |  |  | 3 |  | 0.049998 | 3e-08
19 kΩ | 18999.631 | 18999.626 | 9 |  |  |  | 3 |  | 0.017544 | 2e-08
100 kΩ | 99998.19 | 99998.16 | 11 |  |  |  | 5 |  | 0.050001 | 2e-08
190 kΩ | 189993.6 | 189993.5 | 11 |  |  |  | 5 |  | 0.079436 | 2e-08
1 MΩ | 999904.9 | 999899.9 | 16 |  |  |  | 8 |  | 0.19367 | 4e-08
1.9 MΩ | 1899846.3 | 1899825.7 | 17 |  |  |  | 8 |  | 0.031631 | 2e-08
10MΩ | 9998563 | 9998328 | 33 |  |  |  | 12 |  | 0.13336 | 1.3e-07
19 MΩ | 18999517 | 18999681 | 43 |  |  |  | 12 |  | 0.24966 | 7e-08
100MΩ(2W) | 100007650 | 100008380 | 110 |  |  |  | 180 |  | 1.4857 | 1.7e-07
OHM Zero 4W | DUT Range | Measured ohm | Source Unc. ohm | Lower Limit ohm | Upper Limit ohm | Measured | DUT Transfer STB ppm | Test Result % of SPEC | MEAS SDEV ohm
0 Ω | 10 Ω | -0.000113 | 5e-05 |  |  | N/A |  |  | 1.201e-05 | 1.721e-06
0 Ω | 100 Ω | 0.0002 | 5e-05 |  |  | N/A |  |  | 5e-06 | 2.57e-06
0 Ω | 1 KΩ | 0.0021 | 5e-05 |  |  | N/A |  |  | 0 | 2.2e-05
0 Ω | 10 KΩ | 0.022 | 5e-05 |  |  | N/A |  |  | 0 | 0.00025
0 Ω | 100 KΩ | 0.05 | 5e-05 |  |  | N/A |  |  | 0 | 0.000609
0 Ω | 1 MΩ | 0.5 | 5e-05 |  |  | N/A |  |  | 0 | 0.00686
0 Ω | 10 MΩ | 4 | 5e-05 |  |  | N/A |  |  | 0.5 | 0.0551
0 Ω | 100 MΩ(2W) | 40 | 5e-05 |  |  | N/A |  |  | 0 | 0.527
OHM Zero 2W | DUT Range | Measured ohm | Source Unc. ohm | Lower Limit ohm | Upper Limit ohm | Measured | DUT Transfer STB ppm | Test Result % of SPEC | MEAS SDEV ohm
0 Ω | 10 Ω | 0.26763 | 5e-05 |  |  | N/A |  |  | 0.00014104 | 1.722e-06
0 Ω | 100 Ω | 0.26398 | 5e-05 |  |  | N/A |  |  | 4.7726e-05 | 2.57e-06
0 Ω | 1 KΩ | 0.2658 | 5e-05 |  |  | N/A |  |  | 7.8174e-05 | 2.5e-05
0 Ω | 10 KΩ | 0.286 | 5e-05 |  |  | N/A |  |  | 0.00052705 | 0.000261
0 Ω | 100 KΩ | 0.31 | 5e-05 |  |  | N/A |  |  | 0 | 0.000817
0 Ω | 1 MΩ | 0.7 | 5e-05 |  |  | N/A |  |  | 0.05 | 0.00827
0 Ω | 10 MΩ | 5 | 5e-05 |  |  | N/A |  |  | 0 | 0.0462
0 Ω | 100 MΩ | 40 | 5e-05 |  |  | N/A |  |  | 0 | 0.512
ACV PERFORMANCE TEST
ACV Linearity Checks | DUT Range | Measured Vac | Source Unc. ppm | Lower Limit Vac | Upper Limit Vac | Measured ppm of Input | DUT Transfer STB ppm | Test Result % of SPEC | MEAS SDEV ppm
1VAC/1kHz | 10 VAC | 0.9992 | 77 |  |  |  | 10 |  | 36.4297 | 3.82e-06
2VAC/1kHz | 10 VAC | 1.9995 | 73.5 |  |  |  | 10 |  | 3.9096 | 1.67e-06
5VAC/1kHz | 10 VAC | 4.9999 | 84 |  |  |  | 10 |  | 1.0541 | 1.25e-06
10VAC/1kHz | 10 VAC | 10 | 77 |  |  |  | 10 |  | 0.83333 | 1.09e-06
12VAC/1kHz | 10 VAC | 12.0001 | 75.833 |  |  |  | 10 |  | 0.93169 | 1.32e-06
15VAC/1kHz | 10 VAC | 15.0002 | 74.667 |  |  |  | 10 |  | 0.81649 | 1.42e-06
19VAC/1kHz | 10 VAC | 19.0003 | 73.684 |  |  |  | 10 |  | 0.83216 | 1.2e-06
ACV Performance Tset |  | Measured Vac | Source Unc. ppm | Lower Limit Vac | Upper Limit Vac | Measured ppm | DUT Transfer STB ppm | Test Result % of SPEC | MEAS SDEV ppm
0.001 V, 10 Hz |  | 0.00099492 | 5500 |  |  |  | 20 |  | 58.3433 | 6.52e-06
0.001 V, 20 Hz |  | 0.00099508 | 5200 |  |  |  | 20 |  | 15.8896 | 4.38e-06
0.001 V, 30 Hz |  | 0.00099515 | 5200 |  |  |  | 20 |  | 7.1055 | 2.77e-06
0.001 V, 40 Hz |  | 0.00099513 | 5100 |  |  |  | 20 |  | 8.3741 | 4.85e-06
0.001 V, 55 Hz |  | 0.00099545 | 5100 |  |  |  | 20 |  | 6.0367 | 2.86e-06
0.001 V, 300 Hz |  | 0.00099541 | 5100 |  |  |  | 20 |  | 15.8843 | 2.63e-05
0.001 V, 1 KHz |  | 0.00099549 | 5100 |  |  |  | 20 |  | 4.4296 | 9.34e-06
0.001 V, 10 KHz |  | 0.00099551 | 5100 |  |  |  | 20 |  | 7.8526 | 7.47e-06
0.001 V, 20 KHz |  | 0.000996 | 5100 |  |  |  | 20 |  | 8.8546 | 2.13e-05
0.001 V, 30 KHz |  | 0.00099594 | 5340 |  |  |  | 20 |  | 6.6938 | 8.4e-06
0.001 V, 50 KHz |  | 0.00099632 | 5340 |  |  |  | 20 |  | 7.0972 | 5.51e-06
0.001 V, 100 KHz |  | 0.0009991 | 8800 |  |  |  | 50 |  | 7.2714 | 3.77e-06
0.001 V, 200 KHz |  | 0.00099648 | 16100 |  |  |  | 100 |  | 11.2198 | 4.89e-06
0.001 V, 300 KHz |  | 0.00099383 | 16100 |  |  |  | 100 |  | 6.7081 | 7.46e-06
0.001 V, 500 KHz |  | 0.0010019 | 31500 |  |  |  | 200 |  | 5.9976 | 2.03e-05
0.001 V, 1 MHz |  | 0.0010099 | 33000 |  |  |  | 300 |  | 19.4569 | 7.44e-06
0.01 V, 10 Hz |  | 0.0099971 | 1100 |  |  |  | 20 |  | 1.8339 | 3.26e-06
0.01 V, 20 Hz |  | 0.0099968 | 800 |  |  |  | 20 |  | 0.78198 | 1.18e-06
0.01 V, 30 Hz |  | 0.0099967 | 800 |  |  |  | 20 |  | 0.60112 | 1.71e-06
0.01 V, 40 Hz |  | 0.0099964 | 700 |  |  |  | 20 |  | 2.2369 | 2.7e-06
0.01 V, 55 Hz |  | 0.0099964 | 700 |  |  |  | 20 |  | 2.7396 | 2.42e-06
0.01 V, 300 Hz |  | 0.0099963 | 700 |  |  |  | 20 |  | 3.1236 | 3.33e-06
0.01 V, 1 KHz |  | 0.0099964 | 700 |  |  |  | 20 |  | 1.1671 | 1.87e-06
0.01 V, 10 KHz |  | 0.0099975 | 700 |  |  |  | 20 |  | 1.3646 | 9.3e-07
0.01 V, 20 KHz |  | 0.0099985 | 700 |  |  |  | 20 |  | 1.8785 | 2.3e-06
0.01 V, 30 KHz |  | 0.0099989 | 940 |  |  |  | 20 |  | 1.1182 | 7.8e-07
0.01 V, 50 KHz |  | 0.01 | 940 |  |  |  | 20 |  | 1.0541 | 5.7e-07
0.01 V, 100 KHz |  | 0.010005 | 1600 |  |  |  | 50 |  | 0.78137 | 6.5e-07
0.01 V, 200 KHz |  | 0.0099974 | 2600 |  |  |  | 100 |  | 2.683 | 4.3e-07
0.01 V, 300 KHz |  | 0.0099947 | 2600 |  |  |  | 100 |  | 1.0547 | 4.4e-07
0.01 V, 500 KHz |  | 0.010006 | 4500 |  |  |  | 200 |  | 0.88137 | 1.53e-06
0.01 V, 1 MHz |  | 0.010056 | 6000 |  |  |  | 300 |  | 1.2622 | 6.3e-07
0.1 V, 10 Hz |  | 0.10002 | 660 |  |  |  | 20 |  | 10.0482 | 1.024e-05
0.1 V, 20 Hz |  | 0.1 | 300 |  |  |  | 20 |  | 0.99996 | 1.52e-06
0.1 V, 30 Hz |  | 0.1 | 300 |  |  |  | 20 |  | 0.9718 | 2.72e-06
0.1 V, 40 Hz |  | 0.1 | 195 |  |  |  | 20 |  | 0.8819 | 1.45e-06
0.1 V, 55 Hz |  | 0.1 | 195 |  |  |  | 20 |  | 1.4142 | 1.87e-06
0.1 V, 300 Hz |  | 0.1 | 195 |  |  |  | 20 |  | 2.3979 | 1.64e-06
0.1 V, 1 KHz |  | 0.1 | 195 |  |  |  | 20 |  | 1.0138 | 8.6e-07
0.1 V, 10 KHz |  | 0.1 | 195 |  |  |  | 20 |  | 0.66666 | 3.5e-07
0.1 V, 20 KHz |  | 0.1 | 195 |  |  |  | 20 |  | 0.83333 | 5.7e-07
0.1 V, 30 KHz |  | 0.1 | 400 |  |  |  | 20 |  | 0.78173 | 3.9e-07
0.1 V, 50 KHz |  | 0.1 | 400 |  |  |  | 20 |  | 1.0929 | 5.1e-07
0.1 V, 100 KHz |  | 0.1 | 1050 |  |  |  | 50 |  | 0.99999 | 1.1e-06
0.1 V, 200 KHz |  | 0.10002 | 1240 |  |  |  | 100 |  | 0.70696 | 4.6e-07
0.1 V, 300 KHz |  | 0.099992 | 1240 |  |  |  | 100 |  | 0.72654 | 5.3e-07
0.1 V, 500 KHz |  | 0.099958 | 1900 |  |  |  | 200 |  | 0.83368 | 6e-07
0.1 V, 1 MHz |  | 0.10006 | 4000 |  |  |  | 300 |  | 1.715 | 7.4e-07
1 V, 10 Hz |  | 1.0001 | 600 |  |  |  | 10 |  | 11.8414 | 9.36e-06
1 V, 20 Hz |  | 1 | 180 |  |  |  | 10 |  | 1.1304 | 2.05e-06
1 V, 30 Hz |  | 1 | 180 |  |  |  | 10 |  | 1.0138 | 2.7e-06
1 V, 40 Hz |  | 0.99999 | 77 |  |  |  | 10 |  | 1.118 | 1.38e-06
1 V, 55 Hz |  | 0.99999 | 77 |  |  |  | 10 |  | 0.70712 | 2.18e-06
1 V, 300 Hz |  | 0.99999 | 77 |  |  |  | 10 |  | 1.965 | 1.79e-06
1 V, 1 KHz |  | 1 | 77 |  |  |  | 10 |  | 1.5 | 1.36e-06
1 V, 10 KHz |  | 0.99999 | 77 |  |  |  | 10 |  | 0.97184 | 3.1e-07
1 V, 20 KHz |  | 1 | 77 |  |  |  | 10 |  | 0.8819 | 3.4e-07
1 V, 30 KHz |  | 1 | 140 |  |  |  | 10 |  | 0.866 | 6.1e-07
1 V, 50 KHz |  | 1.0001 | 140 |  |  |  | 20 |  | 0.92791 | 4.8e-07
1 V, 100 KHz |  | 0.9998 | 310 |  |  |  | 30 |  | 1.2363 | 6.4e-07
1 V, 200 KHz |  | 0.99937 | 550 |  |  |  | 70 |  | 0.60131 | 3.4e-07
1 V, 300 KHz |  | 0.99964 | 550 |  |  |  | 70 |  | 0.78201 | 4.4e-07
1 V, 500 KHz |  | 0.99897 | 1400 |  |  |  | 100 |  | 0.92891 | 6e-07
1 V, 1 MHz |  | 0.99954 | 3000 |  |  |  | 200 |  | 0.70743 | 1.59e-06
10 V, 10 Hz |  | 10.0011 | 600 |  |  |  | 10 |  | 6.8549 | 6.95e-06
10 V, 20 Hz |  | 10.0002 | 180 |  |  |  | 10 |  | 1.2018 | 2e-06
10 V, 30 Hz |  | 10.0001 | 180 |  |  |  | 10 |  | 1.0138 | 2.54e-06
10 V, 40 Hz |  | 10 | 77 |  |  |  | 10 |  | 0.86603 | 2.12e-06
10 V, 55 Hz |  | 9.9999 | 77 |  |  |  | 10 |  | 1.453 | 1.98e-06
10 V, 300 Hz |  | 9.9999 | 77 |  |  |  | 10 |  | 2.8771 | 2.27e-06
10 V, 1 KHz |  | 10.0001 | 77 |  |  |  | 10 |  | 1.0929 | 1.03e-06
10 V, 10 KHz |  | 9.9999 | 77 |  |  |  | 10 |  | 0.50001 | 2.8e-07
10 V, 20 KHz |  | 10.0002 | 77 |  |  |  | 10 |  | 0.8819 | 2.4e-07
10 V, 30 KHz |  | 10.0002 | 140 |  |  |  | 10 |  | 1.2018 | 3.9e-07
10 V, 50 KHz |  | 10.0004 | 140 |  |  |  | 20 |  | 0.66664 | 6.1e-07
10 V, 100 KHz |  | 9.9981 | 270 |  |  |  | 30 |  | 1.4243 | 3.2e-07
10 V, 200 KHz |  | 9.9931 | 670 |  |  |  | 70 |  | 0.97249 | 4.6e-07
10 V, 300 KHz |  | 9.9961 | 670 |  |  |  | 70 |  | 1.1185 | 3.7e-07
10 V, 500 KHz |  | 9.9892 | 1700 |  |  |  | 100 |  | 0.50054 | 7.1e-07
10 V, 1 MHz |  | 10.0018 | 3500 |  |  |  | 200 |  | 3.0726 | 3.26e-06
19 V, 10 Hz |  | 19.0067 | 552.63 |  |  |  | 10 |  | 6.8537 | 6.4e-06
19 V, 20 Hz |  | 19.0018 | 165.79 |  |  |  | 10 |  | 1.5116 | 1.39e-06
19 V, 30 Hz |  | 19.0009 | 165.79 |  |  |  | 10 |  | 1.4519 | 2.36e-06
19 V, 40 Hz |  | 19.0005 | 73.684 |  |  |  | 10 |  | 0.66803 | 9.4e-07
19 V, 55 Hz |  | 19.0003 | 73.684 |  |  |  | 10 |  | 1.3187 | 1.82e-06
19 V, 300 Hz |  | 19 | 73.684 |  |  |  | 10 |  | 2.6184 | 1.96e-06
19 V, 1 KHz |  | 19.0001 | 73.684 |  |  |  | 10 |  | 0.90312 | 9.6e-07
19 V, 10 KHz |  | 19 | 73.684 |  |  |  | 10 |  | 0.57521 | 4.7e-07
19 V, 20 KHz |  | 19.0005 | 73.684 |  |  |  | 10 |  | 0.53356 | 4.7e-07
19 V, 30 KHz |  | 19.0007 | 130.53 |  |  |  | 10 |  | 0.51147 | 6e-07
19 V, 50 KHz |  | 19.001 | 130.53 |  |  |  | 20 |  | 0.51146 | 1.8e-07
19 V, 100 KHz |  | 18.9969 | 251.05 |  |  |  | 30 |  | 0.70187 | 3.9e-07
19 V, 200 KHz |  | 18.9878 | 589.47 |  |  |  | 70 |  | 0.68555 | 3.1e-07
19 V, 300 KHz |  | 18.9934 | 589.47 |  |  |  | 70 |  | 0.48857 | 4.1e-07
19 V, 500 KHz |  | 18.9806 | 1463.2 |  |  |  | 100 |  | 0.7655 | 3.8e-07
19 V, 1 MHz |  | 19.0083 | 3073.7 |  |  |  | 200 |  | 9.4651 | 8.7e-07
100 V, 10 Hz |  | 100.0112 | 600 |  |  |  | 10 |  | 5.1256 | 9.94e-06
100 V, 20 Hz |  | 100.0015 | 180 |  |  |  | 10 |  | 1.3017 | 1.03e-06
100 V, 30 Hz |  | 100 | 180 |  |  |  | 10 |  | 0.88192 | 2.31e-06
100 V, 40 Hz |  | 99.9989 | 85 |  |  |  | 10 |  | 0.78174 | 1.52e-06
100 V, 55 Hz |  | 99.9989 | 85 |  |  |  | 10 |  | 0.86603 | 1.97e-06
100 V, 300 Hz |  | 99.9981 | 85 |  |  |  | 10 |  | 2.2913 | 2.52e-06
100 V, 1 KHz |  | 99.9978 | 85 |  |  |  | 10 |  | 1.5812 | 1.13e-06
100 V, 10 KHz |  | 99.998 | 85 |  |  |  | 10 |  | 0.52706 | 1.9e-07
100 V, 20 KHz |  | 99.9972 | 85 |  |  |  | 10 |  | 0.60094 | 5.3e-07
100 V, 30 KHz |  | 99.9979 | 240 |  |  |  | 10 |  | 0.70712 | 3.9e-07
100 V, 50 KHz |  | 99.9966 | 240 |  |  |  | 20 |  | 1.1667 | 5.4e-07
100 V, 100 KHz |  | 99.9906 | 600 |  |  |  | 30 |  | 2.0617 | 1.7e-06
100 V, 200 KHz |  | 99.9853 | 2600 |  |  |  | 50 |  | 3.0872 | 5.34e-06
1000 V, 55 Hz |  | 1000.016 | 79 |  |  |  | 15 |  | 1.3017 | 2.24e-06
1000 V, 300 Hz |  | 1000.024 | 79 |  |  |  | 15 |  | 2.0883 | 1.83e-06
700 V, 1 KHz |  | 699.996 | 80.714 |  |  |  | 50 |  | 1.6148 | 9.9e-07
DCI PERFORMANCE TEST
DCI Test | Expected Value | Measured A | Source Unc. ppm | Lower Limit A | Upper Limit A | Measured ppm | DUT Transfer STB ppm | Test Result % of SPEC | MEAS SDEV ppm
1 µADC | 1e-06 | 9.986e-07 | 10059.0827 | [5 empty] | 50.0701 | 1.126e-05
10 µADC | 1e-05 | 9.999e-06 | 1045.1045 | [5 empty] | 0 | 7.9e-07
50 µADC | 5e-05 | 5.0001e-05 | 245 | [5 empty] | 0.99999 | 2.9e-07
100 µADC | 0.0001 | 0.0001 | 145 |  |  |  | 7 |  | 0.57733 | 1.4e-07
-100 µADC | -0.0001 | -0.00010001 | 145 |  |  |  | 7 |  | -0.49997 | 1.4e-07
-50 µADC | -5e-05 | -5.0003e-05 | 245 | [5 empty] | 0 | 2.6e-07
-10 µADC | -1e-05 | -1.0002e-05 | -1044.8224 | [5 empty] | -4.9992 | 9.5e-07
0.1 mADC | 0.0001 | 9.9995e-05 | 145 | [5 empty] | 5.0003 | 1.28e-06
0.5 mADC | 0.0005 | 0.00050001 | 65 | [5 empty] | 0.99999 | 2e-07
1 mADC | 0.001 | 0.001 | 55 |  |  |  | 7 |  | 0.49999 | 1.2e-07
-1 mADC | -0.001 | -0.001 | 55 |  |  |  | 7 |  | 0 | 1.1e-07
-0.5 mADC | -0.0005 | -0.00050002 | 65 | [5 empty] | -0.99996 | 2.3e-07
1 mADC | 0.001 | 0.001 | 55 | [5 empty] | 5 | 1.25e-06
5 mADC | 0.005 | 0.005 | 65 | [5 empty] | 1 | 2.4e-07
10 mADC | 0.01 | 0.01 | 55 |  |  |  | 7 |  | 0 | 1.3e-07
-10 mADC | -0.01 | -0.01 | 55 |  |  |  | 7 |  | -0.5 | 1e-07
-5 mADC | -0.005 | -0.005 | 65 | [5 empty] | -1 | 2.5e-07
10 mADC | 0.01 | 0.0099983 | 55 | [5 empty] | 0 | 9.7e-07
50 mADC | 0.05 | 0.049999 | 75 | [5 empty] | 0 | 2.3e-07
100 mADC | 0.1 | 0.099998 | 65 |  |  |  | 7 |  | 0.50001 | 1.3e-07
-100 mADC | -0.1 | -0.1 | 65 |  |  |  | 7 |  | -0.49999 | 9e-08
-50 mADC | -0.05 | -0.050001 | 75 | [5 empty] | 0 | 2.4e-07
0.1 ADC | 0.1 | 0.099997 | 65 | [5 empty] | 0 | 1.98e-06
0.5 ADC | 0.5 | 0.49998 | 135 | [5 empty] | 0 | 2.2e-07
-0.5 ADC | -0.5 | -0.50001 | 135 | [5 empty] | -0.99998 | 2.7e-07
-1.0 ADC | -1 | -0.99999 | 105 |  |  |  | 20 |  | -0.5 | 9e-08
1.0 ADC | 1 | 0.99996 | 105 |  |  |  | 20 |  | 0 | 1e-07
ACI PERFORMANCE TEST
ACI Test | Expected Value | Measured Aac | Source Unc. ppm | Lower Limit Aac | Upper Limit Aac | Measured ppm | DUT Transfer STB ppm | Test Result % of SPEC | MEAS SDEV ppm
10 µA, 10 Hz | 1e-05 | 9.9962e-06 | 3651.4058 | [5 empty] | 12.9149 | 1.307e-05
10 µA, 20 Hz | 1e-05 | 9.9959e-06 | 2851.1832 | [5 empty] | 28.7347 | 1.923e-05
10 µA, 30 Hz | 1e-05 | 9.9959e-06 | 2851.169 | [5 empty] | 50.0205 | 3.29e-06
10 µA, 40 Hz | 1e-05 | 9.9959e-06 | 2120.8696 | [5 empty] | 5.0021 | 3.26e-06
10 µA 55 Hz | 1e-05 | 9.996e-06 | 2120.8483 | [5 empty] | 5.002 | 3.82e-06
10 µA, 300 Hz | 1e-05 | 9.9962e-06 | 2120.8165 |  |  | -40 |  |  | 9.578 | 5.07e-06
10 µA, 1 KHz | 1e-05 | 9.9961e-06 | 2120.8271 | [5 empty] | 5.002 | 1.976e-05
10 µA, 5 KHz | 1e-05 | 9.9951e-06 | 5502.7238 | [5 empty] | 14.1491 | 4.31e-06
10 µA, 10 KHz | 1e-05 | 9.9934e-06 | 11507.595 | [5 empty] | 5.0033 | 3.69e-06
100 µA, 10 Hz | 0.0001 | 0.00010002 | 950 | 9.9939e-05 | 0.00010006 | -40 | 50 |  | 2.38 | 1.178e-05
100 µA, 20 Hz | 0.0001 | 0.00010001 | 600 |  |  |  | 50 |  | 2.0614 | 1.86e-06
100 µA, 30 Hz | 0.0001 | 0.00010001 | 600 | 9.9939e-05 | 0.00010006 |  | 50 |  | 0.95736 | 2.54e-06
100 µA, 40 Hz | 0.0001 | 0.00010001 | 320 |  |  |  | 50 |  | 0.99993 | 2.76e-06
100 µA 55 Hz | 0.0001 | 0.00010001 | 320 | 9.9939e-05 | 0.00010006 | -40 | 50 |  | 1.4999 | 2.35e-06
100 µA, 300 Hz | 0.0001 | 0.0001 | 320 |  |  |  | 50 |  | 1.9148 | 1.61e-06
100 µA, 1 KHz | 0.0001 | 0.0001 | 320 | 9.9939e-05 | 0.00010006 |  | 50 |  | 1.4142 | 2.1e-06
100 µA, 5 KHz | 0.0001 | 0.0001 | 0.1 |  |  |  | 100 |  | 0 | 3.9e-07
100 µA, 10 KHz | 0.0001 | 0.00010002 | 0.25 | 9.9939e-05 | 0.00010006 |  | 300 |  | 0.49988 | 7.9e-07
1mA, 10 Hz | 0.001 | 0.0010002 | 700 |  |  |  | 40 |  | 2.8863 | 8.74e-06
1 mA, 20 Hz | 0.001 | 0.0010001 | 390 | 9.9939e-05 | 0.00010006 |  | 40 |  | 0.57731 | 1.4e-06
1 mA, 30 Hz | 0.001 | 0.0010001 | 390 |  |  |  | 40 |  | 0.95738 | 2.98e-06
1 mA, 40 Hz | 0.001 | 0.001 | 160 | 9.9939e-05 | 0.00010006 |  | 40 |  | 0.99996 | 1.98e-06
1 mA, 55 Hz | 0.001 | 0.001 | 160 |  |  |  | 40 |  | 0.95738 | 2.09e-06
1 mA, 300 Hz | 0.001 | 0.001 | 160 | 9.9939e-05 | 0.00010006 |  | 40 |  | 2.5819 | 1.88e-06
1 mA, 1 KHz | 0.001 | 0.001 | 160 |  |  |  | 40 |  | 1.5 | 1.13e-06
1 mA, 5 KHz | 0.001 | 0.001 | 1000 | 9.9939e-05 | 0.00010006 |  | 70 |  | 0.57734 | 5.4e-07
1 mA, 10 KHz | 0.001 | 0.00099999 | 0.25 |  |  |  | 200 |  | 0.50001 | 4.1e-07
10 mA, 10 Hz | 0.01 | 0.010002 | 700 | 9.9939e-05 | 0.00010006 |  | 40 |  | 5.915 | 7.55e-06
10 mA, 20 Hz | 0.01 | 0.010001 | 390 |  |  |  | 40 |  | 0.81644 | 1.53e-06
10 mA, 30 Hz | 0.01 | 0.010001 | 390 | 9.9939e-05 | 0.00010006 |  | 40 |  | 0.57732 | 2.98e-06
10 mA, 40 Hz | 0.01 | 0.01 | 160 |  |  |  | 40 |  | 0.57732 | 2.04e-06
10 mA, 55 Hz | 0.01 | 0.01 | 160 | 9.9939e-05 | 0.00010006 |  | 40 |  | 0.81646 | 2.33e-06
10 mA, 300 Hz | 0.01 | 0.01 | 160 |  |  |  | 40 |  | 2.6298 | 1.76e-06
10 mA, 1 KHz | 0.01 | 0.010001 | 160 | 9.9939e-05 | 0.00010006 |  | 40 |  | 0.95736 | 1.12e-06
10 mA, 5 KHz | 0.01 | 0.010001 | 1000 |  |  |  | 70 |  | 0 | 4.4e-07
10 mA, 10 KHz | 0.01 | 0.0099999 | 0.25 | 9.9939e-05 | 0.00010006 |  | 200 |  | 0.57736 | 3.9e-07
100 mA, 10 Hz | 0.1 | 0.10002 | 700 |  |  |  | 40 |  | 4.9906 | 7.98e-06
100 mA, 20 Hz | 0.1 | 0.10001 | 390 | 9.9939e-05 | 0.00010006 |  | 40 |  | 0.81641 | 1.69e-06
100 mA, 30 Hz | 0.1 | 0.10001 | 390 |  |  |  | 40 |  | 0.81642 | 2.36e-06
100 mA, 40 Hz | 0.1 | 0.10001 | 160 | 9.9939e-05 | 0.00010006 |  | 40 |  | 1.4141 | 1.79e-06
100 mA, 55 Hz | 0.1 | 0.10001 | 160 |  |  |  | 40 |  | 1.9147 | 2.14e-06
100 mA, 300 Hz | 0.1 | 0.10001 | 160 | 9.9939e-05 | 0.00010006 |  | 40 |  | 0.81643 | 2.81e-06
100 mA, 1 KHz | 0.1 | 0.10001 | 160 |  |  |  | 40 |  | 0.95734 | 1.09e-06
100 mA, 5 KHz | 0.1 | 0.099997 | 0.1 | 9.9939e-05 | 0.00010006 |  | 70 |  | 1.2583 | 8.1e-07
100 mA, 10 KHz | 0.1 | 0.099991 | 0.25 |  |  |  | 200 |  | 0.5774 | 8.1e-07
1 A, 10 Hz | 1 | 0.99997 | -1 | 9.9939e-05 | 0.00010006 |  | 40 |  | 7.9375 | 8.24e-06
1 A, 20 Hz | 1 | 0.99997 | -1 |  |  |  | 40 |  | 1.4143 | 9.5e-07
1 A, 30 Hz | 1 | 0.99997 | -1 | 9.9939e-05 | 0.00010006 |  | 40 |  | 0.50002 | 3.4e-06
1 A, 40 Hz | 1 | 1.0001 | 640 |  |  |  | 40 |  | 1.2582 | 1.77e-06
1 A, 55 Hz | 1 | 1.0001 | 640 | 9.9939e-05 | 0.00010006 |  | 40 |  | 1.8928 | 2.4e-06
1 A, 300 Hz | 1 | 1.0001 | 640 |  |  |  | 40 |  | 2.8721 | 2.61e-06
1 A, 1 KHz | 1 | 1 | 640 | 9.9939e-05 | 0.00010006 |  | 40 |  | 0.57733 | 1.04e-06
1 A, 5 KHz | 1 | 0.99942 | 800 |  |  |  | 70 |  | 1.2917 | 8.7e-07
1 A, 10 KHz | 1 | 1 | 0.82 | 9.9939e-05 | 0.00010006 |  | 200 |  | 1.291 | 4.9e-07
Test Summary | Initialization | DCV | OHM | ACV | DCI | ACI |  | Total
Time escape minute | 0.19 | 206.92 | 249.15 | 778.96 | 40.4 | 185.57