## Sheet
PERFORMANCE TEST
DUT
Manufacturer |  | Datron Instruments |  |  | S/N |  | 31255
Model Number |  | 4950 |  |  | Test date |  | 23-09-29 13:09:22
Meter Info |  | Datron 4950 STD |  |  | Ambient Temp. |  | 23.0
Options |  | 1,1,1 |  |  | CERT_AMB? |  | 23
Firmware |  | 890197-02.07 |  |  | TEMP? |  | 41.99
Date of Baseline Cal |  | "1999-SEP-16" |  |  | TEMP? BASE |  | 1.99
Date of Certified Cal |  | "2023-SEP-24" |  |  | TEMP? CERTIFIED |  | -1.2
CNFTST? |  | 0 |  |  | STANDBY? |  | "03:09"
DUT Confidence level |  | 30 Days Transfer Satbility
Calibrator
Reference standard |  | Options |  | Unc | Calibration Date |  | Due Date
Fluke 5700A |  |  |  | -- | -- |  | --
ADRmu12 |  | 10.000000 VDC |  | -- | -- |  | --
Fluke 742A-1 8927002 |  | 1.0000195 ohm |  | -- | -- |  | --
Fluke 742A-10k 7815006 |  | 10.000133 kohm
MFC
MFC Calibrate Date |  | 030622 |  |  | MFC last calibrated |  | 483 days
Calibrate Date WBFLAT |  | 100589 |  |  | MFC since WBFLAT |  | 0 days
Calibrate Date WB Gain |  | 250190 |  |  | MFC since WBGAIN |  | 0 days
MFC Calibrate Date Zero |  | 240923 |  |  | MFC since DCV Zero |  | 5 days
MFC Confidence level |  | 5700A 99% 90 Days |  |  | CAL CONST 11V Range +zero |  | 398.1796
CAL CONST 6.5V |  | +6.529921596265443E+00 |  |  | CAL CONST 11V Range -zero |  | 398.1789
CAL CONST 13V reference voltage |  | +1.306136884003510E+01 |  |  | CAL CONST, Zero calibration temperature |  | 23
CAL CONST 10KOHM true output resistance |  | +1.000039902274076E+04 |  |  | CAL CONST, ALL calibration temp |  | 23
CAL SONST 10kOHM standard resistance |  | +1.000059914344082E+04 |  |  | CAL_CONST? CHECK, WB_FLAT_TEMP |  | 23.9
CAL CONST 1 OHM |  | +9.996206897944441E-01 |  |  | CAL CONST,WBAND_TEMP
ETIME? |  | 6053 days 8.85 hours |  |  | CAL TEMP? CHECK
ONTIME? |  | 0 days 2.85 hours |  |  | CAL CONST,DACLIN |  | 0.359
DCV PERFORMANCE TEST
DCV Test VDC | Expected Value | Measured μV | Source Unc. μV | Lower Limit μV | Upper Limit μV | Deviation | DUT Transfer STB ppm | Test Result % of SPEC | MEAS SDEV μV
0.0mVDC | 0 | 0.05 | 7.5e-07 |  |  | N/A |  |  | 0.031929 | 0.00602
0.0VDC | 0 | 0.9 | 7.5e-07 |  |  | N/A |  |  | 0.05 | 0.01338
0.0VDC | 0 | 7 | 7.5e-07 |  |  | N/A |  |  | 0.33333 | 0.0757
0.0VDC | 0 | 20 | 7.5e-07 |  |  | N/A |  |  | 0 | 0.85
0.0VDC | 0 | 400 | 7.5e-07 |  |  | N/A |  |  | 0 | 3.56
DCV Test VDC | DMM Range | Measured V | Source Unc. ppm | Lower Limit V | Upper Limit V | Measured ppm | DUT Transfer STB ppm | Test Result % of SPEC | MEAS SDEV ppm
0.01 | 100mV | 0.01 | 81.5 | [5 empty] | 0.66665 | 7e-07
0.02 | 100mV | 0.02 | 44 | [5 empty] | 0.25 | 3.4e-07
0.05 | 100mV | 0.05 | 21.5 | [5 empty] | 0.1732 | 1e-07
0.1 | 100mV | 0.1 | 14 |  |  |  | 3 |  | 0.072648 | 6e-08
-0.01 | 100mV | -0.0099998 | -81.5 | [5 empty] | -1.5093 | 5.5e-07
-0.02 | 100mV | -0.02 | -44 | [5 empty] | -0.35356 | 3.5e-07
-0.05 | 100mV | -0.05 | 21.5 | [5 empty] | -0.36209 | 1.1e-07
-0.1 | 100mV | -0.1 | 14 |  |  |  | 3 |  | -0.072648 | 5e-08
0.1 | 1V | 0.1 | 14 | [5 empty] | 1.0541 | 1.3e-07
0.2 | 1V | 0.2 | 10.25 | [5 empty] | 0.25 | 7e-08
0.5 | 1V | 0.5 | 5.9 | [5 empty] | 0.1732 | 4e-08
1 | 1V | 1 | 4.7 |  |  |  | 1.5 |  | 0.11304 | 2e-08
-0.1 | 1V | -0.099999 | 14 | [5 empty] | -0.86603 | 1.4e-07
-0.2 | 1V | -0.2 | 10.25 | [5 empty] | -0.22048 | 6e-08
-0.5 | 1V | -0.5 | 5.9 | [5 empty] | -0.12019 | 3e-08
-1 | 1V | -1 | 4.7 |  |  |  | 1.5 |  | -0.05 | 2e-08
1 | 10V | 1 | 4.7 | [5 empty] | 0.66666 | 8e-08
2 | 10V | 2 | 4.1 | [5 empty] | 0.26352 | 5e-08
5 | 10V | 5 | 4.1 | [5 empty] | 0.14142 | 3e-08
10 | 10V | 10 | 3.8 |  |  |  | 1.5 |  | 0.072648 | 2e-08
19 | 10V | 19 | 3.8158 |  |  |  | 1.5 |  | 0.04386 | 1e-08
-1 | 10V | -0.99999 | 4.7 | [5 empty] | -1.7321 | 8e-08
-2 | 10V | -2 | 4.1 | [5 empty] | -0.16667 | 4e-08
-5 | 10V | -5 | 4.1 | [5 empty] | 0 | 3e-08
-10 | 10V | -10 | 3.8 |  |  |  | 0.5 |  | -0.066667 | 2e-08
-19 | 10V | -19 | 3.8158 |  |  |  | 1.5 |  | -0.10526 | 1e-08
10 | 100V | 10 | 3.8 | [5 empty] | 1.2019 | 9e-08
20 | 100V | 20 | 3.8 | [5 empty] | 0.25 | 5e-08
50 | 100V | 50 | 7 | [5 empty] | 0.1 | 3e-08
100 | 100V | 100 | 6 |  |  |  | 2 |  | 0.1118 | 2e-08
-10 | 100V | -10 | 3.8 | [5 empty] | -0.72648 | 1e-07
-20 | 100V | -20 | 3.8 | [5 empty] | 0 | 4e-08
-50 | 100V | -50 | 7 | [5 empty] | -0.066667 | 3e-08
-100 | 100V | -100 | 6 |  |  |  | 2 |  | -0.070711 | 2e-08
-100 | 1000V | -99.9997 | 6 | [5 empty] | -0.44096 | 4e-08
-200 | 1000V | -199.9995 | 5.5 | [5 empty] | 0 | 3e-08
-500 | 1000V | -500.0008 | 8.2 | [5 empty] | 0 | 2e-08
-1000 | 1000V | -1000.0019 | 7.6 |  |  |  | 2 |  | -0.060092 | 2e-08
100 | 1000V | 100.0003 | 6 | [5 empty] | 1.0541 | 7e-08
200 | 1000V | 200.0004 | 5.5 | [5 empty] | 0.22048 | 3e-08
500 | 1000V | 500.0013 | 8.2 | [5 empty] | 0.1 | 3e-08
1000 | 1000V | 1000.0025 | 7.6 |  |  |  | 2 |  | 0.033333 | 1e-08
OHM PERFORMANCE TEST
OHM Test 4W | Calibrator ohm | Measured ohm | Source Unc. ppm | Lower Limit ohm | Upper Limit ohm | Measured ppm | DUT Transfer STB ppm | Test Result % of SPEC | MEAS SDEV ppm
1 Ω | 0.99962 | 0.99959 | 85 |  |  |  | 20 |  | 2.501 | 1.64e-06
1.9 Ω | 1.9 | 1.8999 | 85 |  |  |  | 15 |  | 0.87723 | 8e-07
10 Ω | 10.0003 | 10.0003 | 26 |  |  |  | 5 |  | 0.099997 | 1.7e-07
19 Ω | 19.0037 | 19.0038 | 24 |  |  |  | 5 |  | 0.08503 | 8e-08
100 Ω | 100.0049 | 100.0053 | 15 |  |  |  | 3 |  | 0.072644 | 3e-08
190 Ω | 190.0021 | 190.0033 | 15 |  |  |  | 3 |  | 0.038235 | 3e-08
1 kΩ | 999.9422 | 999.9455 | 11 |  |  |  | 3 |  | 0.052708 | 3e-08
1.9 kΩ | 1899.9031 | 1899.9052 | 11 |  |  |  | 3 |  | 0.035089 | 2e-08
10 kΩ | 10000.399 | 10000.417 | 9 |  |  |  | 3 |  | 0.06009 | 4e-08
19 kΩ | 18999.631 | 18999.628 | 9 |  |  |  | 3 |  | 0.045581 | 2e-08
100 kΩ | 99998.19 | 99998.17 | 11 |  |  |  | 5 |  | 0.052706 | 2e-08
190 kΩ | 189993.6 | 189993.5 | 11 |  |  |  | 5 |  | 0.07895 | 2e-08
1 MΩ | 999904.9 | 999898.6 | 16 |  |  |  | 8 |  | 0.20618 | 3e-08
1.9 MΩ | 1899846.3 | 1899820 | 17 |  |  |  | 8 |  | 0.045585 | 3e-08
10MΩ | 9998563 | 9998206 | 33 |  |  |  | 12 |  | 0.12695 | 9e-08
19 MΩ | 18999517 | 18999138 | 43 |  |  |  | 12 |  | 0.14414 | 6e-08
100MΩ(2W) | 100007650 | 99996970 | 110 |  |  |  | 180 |  | 1.5739 | 1.6e-07
OHM Zero 4W | DUT Range | Measured ohm | Source Unc. ohm | Lower Limit ohm | Upper Limit ohm | Measured | DUT Transfer STB ppm | Test Result % of SPEC | MEAS SDEV ohm
0 Ω | 10 Ω | -0.000113 | 5e-05 |  |  | N/A |  |  | 1.2845e-05 | 1.6e-06
0 Ω | 100 Ω | 0.00021 | 5e-05 |  |  | N/A |  |  | 5e-06 | 2.8e-06
0 Ω | 1 KΩ | 0.0022 | 5e-05 |  |  | N/A |  |  | 3.3333e-05 | 2.25e-05
0 Ω | 10 KΩ | 0.023 | 5e-05 |  |  | N/A |  |  | 0.00052705 | 0.000243
0 Ω | 100 KΩ | 0.05 | 5e-05 |  |  | N/A |  |  | 0.0044096 | 0.000812
0 Ω | 1 MΩ | 0.5 | 5e-05 |  |  | N/A |  |  | 0 | 0.00734
0 Ω | 10 MΩ | 4 | 5e-05 |  |  | N/A |  |  | 0 | 0.0566
0 Ω | 100 MΩ(2W) | 50 | 5e-05 |  |  | N/A |  |  | 5.2705 | 0.637
OHM Zero 2W | DUT Range | Measured ohm | Source Unc. ohm | Lower Limit ohm | Upper Limit ohm | Measured | DUT Transfer STB ppm | Test Result % of SPEC | MEAS SDEV ohm
0 Ω | 10 Ω | 0.27939 | 5e-05 |  |  | N/A |  |  | 0.00018786 | 2.04e-06
0 Ω | 100 Ω | 0.27565 | 5e-05 |  |  | N/A |  |  | 6.9001e-05 | 2.68e-06
0 Ω | 1 KΩ | 0.2774 | 5e-05 |  |  | N/A |  |  | 7.8174e-05 | 2.19e-05
0 Ω | 10 KΩ | 0.297 | 5e-05 |  |  | N/A |  |  | 0.00044096 | 0.000246
0 Ω | 100 KΩ | 0.33 | 5e-05 |  |  | N/A |  |  | 0.0033333 | 0.001006
0 Ω | 1 MΩ | 0.8 | 5e-05 |  |  | N/A |  |  | 0 | 0.00682
0 Ω | 10 MΩ | 4 | 5e-05 |  |  | N/A |  |  | 0 | 0.0562
0 Ω | 100 MΩ | 40 | 5e-05 |  |  | N/A |  |  | 4.4096 | 0.516
ACV PERFORMANCE TEST
ACV Linearity Checks | DUT Range | Measured Vac | Source Unc. ppm | Lower Limit Vac | Upper Limit Vac | Measured ppm of Input | DUT Transfer STB ppm | Test Result % of SPEC | MEAS SDEV ppm
1VAC/1kHz | 10 VAC | 0.99893 | 77 |  |  |  | 10 |  | 81.7541 | 2.57e-06
2VAC/1kHz | 10 VAC | 1.9994 | 73.5 |  |  |  | 10 |  | 9.5042 | 1.28e-06
5VAC/1kHz | 10 VAC | 4.9998 | 84 |  |  |  | 10 |  | 1.7639 | 1.42e-06
10VAC/1kHz | 10 VAC | 10 | 77 |  |  |  | 10 |  | 1.3017 | 1.04e-06
12VAC/1kHz | 10 VAC | 12.0001 | 75.833 |  |  |  | 10 |  | 1.9982 | 1.08e-06
15VAC/1kHz | 10 VAC | 15.0002 | 74.667 |  |  |  | 10 |  | 0.9428 | 9.1e-07
19VAC/1kHz | 10 VAC | 19.0003 | 73.684 |  |  |  | 10 |  | 0.77965 | 1.32e-06
ACV Performance Tset |  | Measured Vac | Source Unc. ppm | Lower Limit Vac | Upper Limit Vac | Measured ppm | DUT Transfer STB ppm | Test Result % of SPEC | MEAS SDEV ppm
0.001 V, 10 Hz |  | 0.00099487 | 5500 |  |  |  | 20 |  | 69.8204 | 6.98e-06
0.001 V, 20 Hz |  | 0.00099503 | 5200 |  |  |  | 20 |  | 16.7499 | 2.52e-06
0.001 V, 30 Hz |  | 0.0009951 | 5200 |  |  |  | 20 |  | 7.8559 | 3e-06
0.001 V, 40 Hz |  | 0.00099509 | 5100 |  |  |  | 20 |  | 6.0389 | 3.11e-06
0.001 V, 55 Hz |  | 0.00099541 | 5100 |  |  |  | 20 |  | 6.037 | 3.21e-06
0.001 V, 300 Hz |  | 0.00099538 | 5100 |  |  |  | 20 |  | 13.3952 | 2.54e-05
0.001 V, 1 KHz |  | 0.00099544 | 5100 |  |  |  | 20 |  | 10.9791 | 1.026e-05
0.001 V, 10 KHz |  | 0.00099547 | 5100 |  |  |  | 20 |  | 8.3713 | 4.21e-06
0.001 V, 20 KHz |  | 0.00099596 | 5100 |  |  |  | 20 |  | 17.7099 | 2.95e-05
0.001 V, 30 KHz |  | 0.00099591 | 5340 |  |  |  | 20 |  | 8.6958 | 5.22e-06
0.001 V, 50 KHz |  | 0.00099627 | 5340 |  |  |  | 20 |  | 8.6927 | 4.35e-06
0.001 V, 100 KHz |  | 0.00099899 | 8800 |  |  |  | 50 |  | 7.0782 | 3.56e-06
0.001 V, 200 KHz |  | 0.00099638 | 16100 |  |  |  | 100 |  | 11.345 | 5.45e-06
0.001 V, 300 KHz |  | 0.0009938 | 16100 |  |  |  | 100 |  | 6.7083 | 7.7e-06
0.001 V, 500 KHz |  | 0.0010018 | 31500 |  |  |  | 200 |  | 4.9911 | 1.617e-05
0.001 V, 1 MHz |  | 0.0010092 | 33000 |  |  |  | 300 |  | 9.1951 | 9.36e-06
0.01 V, 10 Hz |  | 0.0099978 | 1100 |  |  |  | 20 |  | 1.4817 | 2.37e-06
0.01 V, 20 Hz |  | 0.0099969 | 800 |  |  |  | 20 |  | 0.7267 | 1.2e-06
0.01 V, 30 Hz |  | 0.0099967 | 800 |  |  |  | 20 |  | 1.564 | 1.1e-06
0.01 V, 40 Hz |  | 0.0099964 | 700 |  |  |  | 20 |  | 1.5005 | 1.94e-06
0.01 V, 55 Hz |  | 0.0099965 | 700 |  |  |  | 20 |  | 0.83362 | 1.54e-06
0.01 V, 300 Hz |  | 0.0099964 | 700 |  |  |  | 20 |  | 2.1674 | 3.45e-06
0.01 V, 1 KHz |  | 0.0099965 | 700 |  |  |  | 20 |  | 1.642 | 1.12e-06
0.01 V, 10 KHz |  | 0.0099976 | 700 |  |  |  | 20 |  | 0.70728 | 7.4e-07
0.01 V, 20 KHz |  | 0.0099986 | 700 |  |  |  | 20 |  | 2.3157 | 2.56e-06
0.01 V, 30 KHz |  | 0.0099989 | 940 |  |  |  | 20 |  | 0.88201 | 8.2e-07
0.01 V, 50 KHz |  | 0.01 | 940 |  |  |  | 20 |  | 0.52704 | 8.9e-07
0.01 V, 100 KHz |  | 0.010004 | 1600 |  |  |  | 50 |  | 0.7068 | 5.9e-07
0.01 V, 200 KHz |  | 0.0099969 | 2600 |  |  |  | 100 |  | 1.2364 | 8.8e-07
0.01 V, 300 KHz |  | 0.0099942 | 2600 |  |  |  | 100 |  | 0.78219 | 6.4e-07
0.01 V, 500 KHz |  | 0.010005 | 4500 |  |  |  | 200 |  | 2.832 | 1.51e-06
0.01 V, 1 MHz |  | 0.010052 | 6000 |  |  |  | 300 |  | 1.407 | 1.08e-06
0.1 V, 10 Hz |  | 0.10002 | 660 |  |  |  | 20 |  | 9.0875 | 8.78e-06
0.1 V, 20 Hz |  | 0.1 | 300 |  |  |  | 20 |  | 1.4142 | 7.8e-07
0.1 V, 30 Hz |  | 0.1 | 300 |  |  |  | 20 |  | 0.83331 | 2.07e-06
0.1 V, 40 Hz |  | 0.1 | 195 |  |  |  | 20 |  | 2.6509 | 2.56e-06
0.1 V, 55 Hz |  | 0.1 | 195 |  |  |  | 20 |  | 0.83332 | 2.58e-06
0.1 V, 300 Hz |  | 0.1 | 195 |  |  |  | 20 |  | 1.1304 | 1.49e-06
0.1 V, 1 KHz |  | 0.1 | 195 |  |  |  | 20 |  | 1.5635 | 9.2e-07
0.1 V, 10 KHz |  | 0.1 | 195 |  |  |  | 20 |  | 0.88191 | 5.4e-07
0.1 V, 20 KHz |  | 0.1 | 195 |  |  |  | 20 |  | 0.83333 | 7.3e-07
0.1 V, 30 KHz |  | 0.1 | 400 |  |  |  | 20 |  | 1.118 | 3.7e-07
0.1 V, 50 KHz |  | 0.1 | 400 |  |  |  | 20 |  | 0.44096 | 3.9e-07
0.1 V, 100 KHz |  | 0.099999 | 1050 |  |  |  | 50 |  | 0.72649 | 9e-07
0.1 V, 200 KHz |  | 0.10001 | 1240 |  |  |  | 100 |  | 0.88179 | 6e-07
0.1 V, 300 KHz |  | 0.099985 | 1240 |  |  |  | 100 |  | 0.9281 | 3e-07
0.1 V, 500 KHz |  | 0.099944 | 1900 |  |  |  | 200 |  | 0.88241 | 5.7e-07
0.1 V, 1 MHz |  | 0.10001 | 4000 |  |  |  | 300 |  | 1.118 | 1e-06
1 V, 10 Hz |  | 1.0001 | 600 |  |  |  | 10 |  | 9.4874 | 9.81e-06
1 V, 20 Hz |  | 1 | 180 |  |  |  | 10 |  | 0.99999 | 1.01e-06
1 V, 30 Hz |  | 1 | 180 |  |  |  | 10 |  | 0.83334 | 2.28e-06
1 V, 40 Hz |  | 0.99999 | 77 |  |  |  | 10 |  | 2.2048 | 1.96e-06
1 V, 55 Hz |  | 0.99999 | 77 |  |  |  | 10 |  | 1.2248 | 2.52e-06
1 V, 300 Hz |  | 0.99999 | 77 |  |  |  | 10 |  | 0.88193 | 1.73e-06
1 V, 1 KHz |  | 1 | 77 |  |  |  | 10 |  | 1.2247 | 1.37e-06
1 V, 10 KHz |  | 0.99998 | 77 |  |  |  | 10 |  | 1.3945 | 2.9e-07
1 V, 20 KHz |  | 1 | 77 |  |  |  | 10 |  | 0.49999 | 5.9e-07
1 V, 30 KHz |  | 1 | 140 |  |  |  | 10 |  | 0.49999 | 4.6e-07
1 V, 50 KHz |  | 1 | 140 |  |  |  | 20 |  | 0.6009 | 3.5e-07
1 V, 100 KHz |  | 0.99979 | 310 |  |  |  | 30 |  | 0.8821 | 9.2e-07
1 V, 200 KHz |  | 0.99934 | 550 |  |  |  | 70 |  | 1.0145 | 4.3e-07
1 V, 300 KHz |  | 0.9996 | 550 |  |  |  | 70 |  | 0.70739 | 5.4e-07
1 V, 500 KHz |  | 0.99886 | 1400 |  |  |  | 100 |  | 0.66743 | 3.3e-07
1 V, 1 MHz |  | 0.99909 | 3000 |  |  |  | 200 |  | 1.4543 | 6.8e-07
10 V, 10 Hz |  | 10.0011 | 600 |  |  |  | 10 |  | 9.9208 | 9.62e-06
10 V, 20 Hz |  | 10.0001 | 180 |  |  |  | 10 |  | 1.6666 | 1.58e-06
10 V, 30 Hz |  | 10 | 180 |  |  |  | 10 |  | 0.52705 | 2.18e-06
10 V, 40 Hz |  | 9.9999 | 77 |  |  |  | 10 |  | 2.6194 | 1.56e-06
10 V, 55 Hz |  | 9.9999 | 77 |  |  |  | 10 |  | 1.0138 | 3.09e-06
10 V, 300 Hz |  | 9.9998 | 77 |  |  |  | 10 |  | 1.3229 | 2.44e-06
10 V, 1 KHz |  | 10 | 77 |  |  |  | 10 |  | 1.4142 | 1.07e-06
10 V, 10 KHz |  | 9.9998 | 77 |  |  |  | 10 |  | 0.52706 | 4.4e-07
10 V, 20 KHz |  | 10.0001 | 77 |  |  |  | 10 |  | 1.2247 | 4e-07
10 V, 30 KHz |  | 10.0001 | 140 |  |  |  | 10 |  | 0.60092 | 2.5e-07
10 V, 50 KHz |  | 10.0003 | 140 |  |  |  | 20 |  | 0.49999 | 2.9e-07
10 V, 100 KHz |  | 9.998 | 270 |  |  |  | 30 |  | 0.70725 | 6e-07
10 V, 200 KHz |  | 9.9928 | 670 |  |  |  | 70 |  | 0 | 3.8e-07
10 V, 300 KHz |  | 9.9958 | 670 |  |  |  | 70 |  | 1.2698 | 7.6e-07
10 V, 500 KHz |  | 9.9884 | 1700 |  |  |  | 100 |  | 1.0553 | 5.8e-07
10 V, 1 MHz |  | 9.9985 | 3500 |  |  |  | 200 |  | 2.8288 | 1.48e-06
19 V, 10 Hz |  | 19.0068 | 552.63 |  |  |  | 10 |  | 5.5369 | 4.71e-06
19 V, 20 Hz |  | 19.0017 | 165.79 |  |  |  | 10 |  | 1.1198 | 1.04e-06
19 V, 30 Hz |  | 19.0008 | 165.79 |  |  |  | 10 |  | 0.59492 | 1.97e-06
19 V, 40 Hz |  | 19.0004 | 73.684 |  |  |  | 10 |  | 4.3639 | 2.21e-06
19 V, 55 Hz |  | 19.0002 | 73.684 |  |  |  | 10 |  | 0.65054 | 2.02e-06
19 V, 300 Hz |  | 18.9999 | 73.684 |  |  |  | 10 |  | 1.5667 | 2.14e-06
19 V, 1 KHz |  | 19.0001 | 73.684 |  |  |  | 10 |  | 1.1403 | 1.01e-06
19 V, 10 KHz |  | 18.9999 | 73.684 |  |  |  | 10 |  | 0.41144 | 3e-07
19 V, 20 KHz |  | 19.0004 | 73.684 |  |  |  | 10 |  | 0.37215 | 3.3e-07
19 V, 30 KHz |  | 19.0006 | 130.53 |  |  |  | 10 |  | 0.5752 | 3.8e-07
19 V, 50 KHz |  | 19.001 | 130.53 |  |  |  | 20 |  | 0.43857 | 5.9e-07
19 V, 100 KHz |  | 18.9967 | 251.05 |  |  |  | 30 |  | 0.87735 | 4.5e-07
19 V, 200 KHz |  | 18.9874 | 589.47 |  |  |  | 70 |  | 0.61444 | 4.5e-07
19 V, 300 KHz |  | 18.993 | 589.47 |  |  |  | 70 |  | 0.95323 | 5.8e-07
19 V, 500 KHz |  | 18.9791 | 1463.2 |  |  |  | 100 |  | 0.5554 | 7.9e-07
19 V, 1 MHz |  | 19.001 | 3073.7 |  |  |  | 200 |  | 7.3936 | 1.19e-06
100 V, 10 Hz |  | 100.0116 | 600 |  |  |  | 10 |  | 6.9053 | 1.296e-05
100 V, 20 Hz |  | 100.0012 | 180 |  |  |  | 10 |  | 0.86602 | 6.6e-07
100 V, 30 Hz |  | 99.9998 | 180 |  |  |  | 10 |  | 0.70711 | 2.02e-06
100 V, 40 Hz |  | 99.9983 | 85 |  |  |  | 10 |  | 1.7321 | 2.63e-06
100 V, 55 Hz |  | 99.9987 | 85 |  |  |  | 10 |  | 1 | 2.7e-06
100 V, 300 Hz |  | 99.9982 | 85 |  |  |  | 10 |  | 2.1859 | 1.65e-06
100 V, 1 KHz |  | 99.9976 | 85 |  |  |  | 10 |  | 1.4814 | 1.38e-06
100 V, 10 KHz |  | 99.9978 | 85 |  |  |  | 10 |  | 0.44097 | 4.7e-07
100 V, 20 KHz |  | 99.9972 | 85 |  |  |  | 10 |  | 0.88194 | 2.9e-07
100 V, 30 KHz |  | 99.9978 | 240 |  |  |  | 10 |  | 0.52706 | 3.6e-07
100 V, 50 KHz |  | 99.9964 | 240 |  |  |  | 20 |  | 0.86606 | 4.8e-07
100 V, 100 KHz |  | 99.9901 | 600 |  |  |  | 30 |  | 1.3335 | 4.5e-07
100 V, 200 KHz |  | 99.9836 | 2600 |  |  |  | 50 |  | 2.0279 | 4.57e-06
1000 V, 55 Hz |  | 1000.016 | 79 |  |  |  | 15 |  | 0.97181 | 2.28e-06
1000 V, 300 Hz |  | 1000.023 | 79 |  |  |  | 15 |  | 2.9907 | 1.66e-06
700 V, 1 KHz |  | 699.996 | 80.714 |  |  |  | 50 |  | 1.1905 | 1.46e-06
DCI PERFORMANCE TEST
DCI Test | Expected Value | Measured A | Source Unc. ppm | Lower Limit A | Upper Limit A | Measured ppm | DUT Transfer STB ppm | Test Result % of SPEC | MEAS SDEV ppm
1 µADC | 1e-06 | 9.985e-07 | 10060.0901 | [5 empty] | 50.0751 | 1.126e-05
10 µADC | 1e-05 | 9.9989e-06 | 1045.115 | [5 empty] | 5.0006 | 1.25e-06
50 µADC | 5e-05 | 5e-05 | 245 | [5 empty] | 0 | 2.2e-07
100 µADC | 0.0001 | 0.0001 | 145 |  |  |  | 7 |  | 0.49999 | 1.2e-07
-100 µADC | -0.0001 | -0.00010001 | 145 |  |  |  | 7 |  | -0.49997 | 1.4e-07
-50 µADC | -5e-05 | -5.0003e-05 | 245 | [5 empty] | -0.99993 | 2.4e-07
-10 µADC | -1e-05 | -1.0002e-05 | -1044.8015 | [5 empty] | 0 | 1.32e-06
0.1 mADC | 0.0001 | 9.9994e-05 | 145 | [5 empty] | 0 | 8.1e-07
0.5 mADC | 0.0005 | 0.0005 | 65 | [5 empty] | 0 | 2.6e-07
1 mADC | 0.001 | 0.001 | 55 |  |  |  | 7 |  | 0.57734 | 1.3e-07
-1 mADC | -0.001 | -0.001 | 55 |  |  |  | 7 |  | -0.49999 | 7e-08
-0.5 mADC | -0.0005 | -0.00050002 | 65 | [5 empty] | 0 | 3.1e-07
1 mADC | 0.001 | 0.001 | 55 | [5 empty] | 5.7735 | 1.24e-06
5 mADC | 0.005 | 0.005 | 65 | [5 empty] | 1 | 3e-07
10 mADC | 0.01 | 0.01 | 55 |  |  |  | 7 |  | 0 | 1.6e-07
-10 mADC | -0.01 | -0.01 | 55 |  |  |  | 7 |  | 0 | 1.1e-07
-5 mADC | -0.005 | -0.005 | 65 | [5 empty] | -1.1547 | 2.7e-07
10 mADC | 0.01 | 0.0099982 | 55 | [5 empty] | 5.0009 | 1.09e-06
50 mADC | 0.05 | 0.049998 | 75 | [5 empty] | 0 | 2.3e-07
100 mADC | 0.1 | 0.099998 | 65 |  |  |  | 7 |  | 0 | 1.1e-07
-100 mADC | -0.1 | -0.1 | 65 |  |  |  | 7 |  | -0.49999 | 1.3e-07
-50 mADC | -0.05 | -0.050001 | 75 | [5 empty] | 0 | 2.8e-07
0.1 ADC | 0.1 | 0.099996 | 65 | [5 empty] | 5.0002 | 1.47e-06
0.5 ADC | 0.5 | 0.49998 | 135 | [5 empty] | 0 | 2.1e-07
-0.5 ADC | -0.5 | -0.50001 | 135 | [5 empty] | 0 | 2.4e-07
-1.0 ADC | -1 | -0.99999 | 105 |  |  |  | 20 |  | -0.50001 | 1e-07
1.0 ADC | 1 | 0.99996 | 105 |  |  |  | 20 |  | 0 | 1.3e-07
ACI PERFORMANCE TEST
ACI Test | Expected Value | Measured Aac | Source Unc. ppm | Lower Limit Aac | Upper Limit Aac | Measured ppm | DUT Transfer STB ppm | Test Result % of SPEC | MEAS SDEV ppm
10 µA, 10 Hz | 1e-05 | 9.9974e-06 | 3650.9675 | [5 empty] | 31.0995 | 2.25e-05
10 µA, 20 Hz | 1e-05 | 9.9969e-06 | 2850.8695 | [5 empty] | 5.7753 | 4.79e-06
10 µA, 30 Hz | 1e-05 | 9.9969e-06 | 2850.8838 | [5 empty] | 5.0016 | 4.39e-06
10 µA, 40 Hz | 1e-05 | 9.9969e-06 | 2120.6468 | [5 empty] | 5.7753 | 2.96e-06
10 µA 55 Hz | 1e-05 | 9.9969e-06 | 2120.6574 | [5 empty] | 5.0016 | 3.61e-06
10 µA, 300 Hz | 1e-05 | 9.9971e-06 | 2120.6044 |  |  | -40 |  |  | 5.7751 | 5.3e-06
10 µA, 1 KHz | 1e-05 | 9.997e-06 | 2120.6256 | [5 empty] | 9.5771 | 2.65e-05
10 µA, 5 KHz | 1e-05 | 9.9963e-06 | 5502.0358 | [5 empty] | 5.0019 | 3.91e-06
10 µA, 10 KHz | 1e-05 | 9.9946e-06 | 11506.2134 | [5 empty] | 5.0027 | 4.19e-06
100 µA, 10 Hz | 0.0001 | 0.00010002 | 950 | 9.9939e-05 | 0.00010006 | -40 | 50 |  | 10.5966 | 6.23e-06
100 µA, 20 Hz | 0.0001 | 0.00010001 | 600 |  |  |  | 50 |  | 0.49995 | 9.9e-07
100 µA, 30 Hz | 0.0001 | 0.00010001 | 600 | 9.9939e-05 | 0.00010006 |  | 50 |  | 1.4141 | 1.94e-06
100 µA, 40 Hz | 0.0001 | 0.00010001 | 320 |  |  |  | 50 |  | 2.2172 | 3.22e-06
100 µA 55 Hz | 0.0001 | 0.00010001 | 320 | 9.9939e-05 | 0.00010006 | -40 | 50 |  | 1.2909 | 2.97e-06
100 µA, 300 Hz | 0.0001 | 0.0001 | 320 |  |  |  | 50 |  | 2.0615 | 1.78e-06
100 µA, 1 KHz | 0.0001 | 0.0001 | 320 | 9.9939e-05 | 0.00010006 |  | 50 |  | 0.57734 | 3.4e-06
100 µA, 5 KHz | 0.0001 | 0.0001 | 0.1 |  |  |  | 100 |  | 0.49998 | 7.2e-07
100 µA, 10 KHz | 0.0001 | 0.00010002 | 0.25 | 9.9939e-05 | 0.00010006 |  | 300 |  | 0.81631 | 5.3e-07
1mA, 10 Hz | 0.001 | 0.0010002 | 700 |  |  |  | 40 |  | 5.4991 | 1.12e-05
1 mA, 20 Hz | 0.001 | 0.0010001 | 390 | 9.9939e-05 | 0.00010006 |  | 40 |  | 0.81644 | 1.27e-06
1 mA, 30 Hz | 0.001 | 0.0010001 | 390 |  |  |  | 40 |  | 0.95738 | 1.59e-06
1 mA, 40 Hz | 0.001 | 0.001 | 160 | 9.9939e-05 | 0.00010006 |  | 40 |  | 2.3093 | 2.53e-06
1 mA, 55 Hz | 0.001 | 0.001 | 160 |  |  |  | 40 |  | 0.99995 | 2.49e-06
1 mA, 300 Hz | 0.001 | 0.001 | 160 | 9.9939e-05 | 0.00010006 |  | 40 |  | 2.3804 | 1.95e-06
1 mA, 1 KHz | 0.001 | 0.001 | 160 |  |  |  | 40 |  | 0.49998 | 9e-07
1 mA, 5 KHz | 0.001 | 0.001 | 1000 | 9.9939e-05 | 0.00010006 |  | 70 |  | 0.49999 | 5.6e-07
1 mA, 10 KHz | 0.001 | 0.00099998 | 0.25 |  |  |  | 200 |  | 0.57736 | 5.4e-07
10 mA, 10 Hz | 0.01 | 0.010002 | 700 | 9.9939e-05 | 0.00010006 |  | 40 |  | 8.6153 | 1.031e-05
10 mA, 20 Hz | 0.01 | 0.010001 | 390 |  |  |  | 40 |  | 0.81644 | 1.05e-06
10 mA, 30 Hz | 0.01 | 0.010001 | 390 | 9.9939e-05 | 0.00010006 |  | 40 |  | 1.2582 | 2.3e-06
10 mA, 40 Hz | 0.01 | 0.010001 | 160 |  |  |  | 40 |  | 2.0614 | 3.16e-06
10 mA, 55 Hz | 0.01 | 0.01 | 160 | 9.9939e-05 | 0.00010006 |  | 40 |  | 1.1547 | 2.21e-06
10 mA, 300 Hz | 0.01 | 0.01 | 160 |  |  |  | 40 |  | 2.6456 | 2.75e-06
10 mA, 1 KHz | 0.01 | 0.010001 | 160 | 9.9939e-05 | 0.00010006 |  | 40 |  | 0.81644 | 1.45e-06
10 mA, 5 KHz | 0.01 | 0.010001 | 1000 |  |  |  | 70 |  | 0.81638 | 3.7e-07
10 mA, 10 KHz | 0.01 | 0.0099998 | 0.25 | 9.9939e-05 | 0.00010006 |  | 200 |  | 0.50001 | 3.8e-07
100 mA, 10 Hz | 0.1 | 0.10002 | 700 |  |  |  | 40 |  | 8.4589 | 8.91e-06
100 mA, 20 Hz | 0.1 | 0.10001 | 390 | 9.9939e-05 | 0.00010006 |  | 40 |  | 2.4492 | 1.3e-06
100 mA, 30 Hz | 0.1 | 0.10001 | 390 |  |  |  | 40 |  | 0.49995 | 2.03e-06
100 mA, 40 Hz | 0.1 | 0.10001 | 160 | 9.9939e-05 | 0.00010006 |  | 40 |  | 2.4998 | 2.18e-06
100 mA, 55 Hz | 0.1 | 0.10001 | 160 |  |  |  | 40 |  | 0.49996 | 2e-06
100 mA, 300 Hz | 0.1 | 0.10001 | 160 | 9.9939e-05 | 0.00010006 |  | 40 |  | 2.0815 | 2.04e-06
100 mA, 1 KHz | 0.1 | 0.10001 | 160 |  |  |  | 40 |  | 0.95734 | 1.51e-06
100 mA, 5 KHz | 0.1 | 0.099997 | 0.1 | 9.9939e-05 | 0.00010006 |  | 70 |  | 0.81652 | 5.5e-07
100 mA, 10 KHz | 0.1 | 0.09999 | 0.25 |  |  |  | 200 |  | 0.50005 | 3.3e-07
1 A, 10 Hz | 1 | 0.99996 | -1 | 9.9939e-05 | 0.00010006 |  | 40 |  | 7.6324 | 6.71e-06
1 A, 20 Hz | 1 | 0.99997 | -1 |  |  |  | 40 |  | 1.4143 | 1.46e-06
1 A, 30 Hz | 1 | 0.99997 | -1 | 9.9939e-05 | 0.00010006 |  | 40 |  | 0.81652 | 1.67e-06
1 A, 40 Hz | 1 | 1.0001 | 640 |  |  |  | 40 |  | 4.7954 | 9.4e-07
1 A, 55 Hz | 1 | 1.0001 | 640 | 9.9939e-05 | 0.00010006 |  | 40 |  | 1.7077 | 2.48e-06
1 A, 300 Hz | 1 | 1.0001 | 640 |  |  |  | 40 |  | 2.7536 | 1.65e-06
1 A, 1 KHz | 1 | 1 | 640 | 9.9939e-05 | 0.00010006 |  | 40 |  | 1.1547 | 1.3e-06
1 A, 5 KHz | 1 | 0.99942 | 800 |  |  |  | 70 |  | 0.50029 | 4.5e-07
1 A, 10 KHz | 1 | 1 | 0.82 | 9.9939e-05 | 0.00010006 |  | 200 |  | 1.1547 | 5.8e-07
Test Summary | Initialization | DCV | OHM | ACV | DCI | ACI |  | Total
Time escape minute | 0.19 | 206.93 | 249.16 | 667.21 | 40.4 | 185.74