## Sheet
PERFORMANCE TEST
DUT
Manufacturer |  | Datron Instruments |  |  | S/N |  | 31255
Model Number |  | 4950 |  |  | Test date |  | 23-09-02 11:05:45
Meter Info |  | Datron 4950 STD |  |  | Ambient Temp. |  | 23.0
Options |  | 1,1,1 |  |  | CERT_AMB? |  | 23
Firmware |  | 890197-02.07 |  |  | TEMP? |  | 43.35
Date of Baseline Cal |  | "1999-SEP-16" |  |  | TEMP? BASE |  | 3.35
Date of Certified Cal |  | "2023-SEP-01" |  |  | TEMP? CERTIFIED |  | 0.14
CNFTST? |  | 0 |  |  | STANDBY? |  | "00:00"
DUT Confidence level |  | 30 Days Transfer Satbility
Calibrator
Reference standard |  | Options |  | Unc | Calibration Date |  | Due Date
Fluke 5700A |  |  |  | -- | -- |  | --
ADRmu12 |  | 10.000000 VDC |  | -- | -- |  | --
Fluke 742A-1 8927002 |  | 1.0000195 ohm |  | -- | -- |  | --
Fluke 742A-10k 7815006 |  | 10.000133 kohm
MFC
MFC Calibrate Date |  | 030622 |  |  | MFC last calibrated |  | 456 days
Calibrate Date WBFLAT |  | 100589 |  |  | MFC since WBFLAT |  | 0 days
Calibrate Date WB Gain |  | 250190 |  |  | MFC since WBGAIN |  | 0 days
MFC Calibrate Date Zero |  | 270823 |  |  | MFC since DCV Zero |  | 6 days
MFC Confidence level |  | 5700A 99% 90 Days |  |  | CAL CONST 11V Range +zero |  | 398.1791
CAL CONST 6.5V |  | +6.529921596265443E+00 |  |  | CAL CONST 11V Range -zero |  | 398.1783
CAL CONST 13V reference voltage |  | +1.306136884003510E+01 |  |  | CAL CONST, Zero calibration temperature |  | 23
CAL CONST 10KOHM true output resistance |  | +1.000039902274076E+04 |  |  | CAL CONST, ALL calibration temp |  | 23
CAL SONST 10kOHM standard resistance |  | +1.000059914344082E+04 |  |  | CAL_CONST? CHECK, WB_FLAT_TEMP |  | 23.9
CAL CONST 1 OHM |  | +9.996206897944441E-01 |  |  | CAL CONST,WBAND_TEMP
ETIME? |  | 6042 days 13.53 hours |  |  | CAL TEMP? CHECK
ONTIME? |  | 0 days 12.80 hours |  |  | CAL CONST,DACLIN |  | 0.359
DCV PERFORMANCE TEST
DCV Test VDC | Expected Value | Measured μV | Source Unc. μV | Lower Limit μV | Upper Limit μV | Deviation | DUT Transfer STB ppm | Test Result % of SPEC | MEAS SDEV μV
0.0mVDC | 0 | -0.04 | 7.5e-07 |  |  | N/A |  |  | 0.017401 | 0.00609
0.0VDC | 0 | 0.2 | 7.5e-07 |  |  | N/A |  |  | 0.05 | 0.01438
0.0VDC | 0 | 0 | 7.5e-07 |  |  | N/A |  |  | 0.44096 | 0.0646
0.0VDC | 0 | 40 | 7.5e-07 |  |  | N/A |  |  | 0 | 0.732
0.0VDC | 0 | 100 | 7.5e-07 |  |  | N/A |  |  | 0 | 5.24
DCV Test VDC | DMM Range | Measured V | Source Unc. ppm | Lower Limit V | Upper Limit V | Measured ppm | DUT Transfer STB ppm | Test Result % of SPEC | MEAS SDEV ppm
0.01 | 100mV | 0.01 | 81.5 | [5 empty] | 2.4037 | 5.9e-07
0.02 | 100mV | 0.02 | 44 | [5 empty] | 0.90139 | 2.5e-07
0.05 | 100mV | 0.05 | 21.5 | [5 empty] | 0.59815 | 1.2e-07
0.1 | 100mV | 0.1 | 14 |  |  |  | 3 |  | 0.22048 | 6e-08
-0.01 | 100mV | -0.0099999 | -81.5 | [5 empty] | -1.2248 | 5.6e-07
-0.02 | 100mV | -0.02 | -44 | [5 empty] | -0.39087 | 2.8e-07
-0.05 | 100mV | -0.05 | 21.5 | [5 empty] | -0.22608 | 1.3e-07
-0.1 | 100mV | -0.1 | 14 |  |  |  | 3 |  | -0.12247 | 5e-08
0.1 | 1V | 0.1 | 14 | [5 empty] | 1 | 1.6e-07
0.2 | 1V | 0.2 | 10.25 | [5 empty] | 0.25 | 8e-08
0.5 | 1V | 0.5 | 5.9 | [5 empty] | 0.14142 | 3e-08
1 | 1V | 1 | 4.7 |  |  |  | 1.5 |  | 0.12247 | 2e-08
-0.1 | 1V | -0.099999 | 14 | [5 empty] | -0.78175 | 1.1e-07
-0.2 | 1V | -0.2 | 10.25 | [5 empty] | -0.26353 | 6e-08
-0.5 | 1V | -0.5 | 5.9 | [5 empty] | -0.066667 | 3e-08
-1 | 1V | -1 | 4.7 |  |  |  | 1.5 |  | -0.088192 | 2e-08
1 | 10V | 1 | 4.7 | [5 empty] | 0.72648 | 8e-08
2 | 10V | 2 | 4.1 | [5 empty] | 0.16667 | 5e-08
5 | 10V | 5 | 4.1 | [5 empty] | 0.066667 | 3e-08
10 | 10V | 10 | 3.8 |  |  |  | 1.5 |  | 0.05 | 2e-08
19 | 10V | 19 | 3.8158 |  |  |  | 1.5 |  | 0.066805 | 1e-08
-1 | 10V | -1 | 4.7 | [5 empty] | -1.6667 | 6e-08
-2 | 10V | -2 | 4.1 | [5 empty] | 0 | 5e-08
-5 | 10V | -5 | 4.1 | [5 empty] | -0.10541 | 2e-08
-10 | 10V | -10 | 3.8 |  |  |  | 0.5 |  | -0.072648 | 2e-08
-19 | 10V | -19 | 3.8158 |  |  |  | 1.5 |  | -0.066805 | 2e-08
10 | 100V | 10.0001 | 3.8 | [5 empty] | 1.2018 | 1e-07
20 | 100V | 20.0002 | 3.8 | [5 empty] | 0.26352 | 5e-08
50 | 100V | 50.0004 | 7 | [5 empty] | 0.066666 | 3e-08
100 | 100V | 100.0008 | 6 |  |  |  | 2 |  | 0.078173 | 2e-08
-10 | 100V | -9.9999 | 3.8 | [5 empty] | -0.72649 | 8e-08
-20 | 100V | -19.9998 | 3.8 | [5 empty] | -0.26353 | 4e-08
-50 | 100V | -49.9996 | 7 | [5 empty] | -0.1 | 3e-08
-100 | 100V | -99.9992 | 6 |  |  |  | 2 |  | -0.070711 | 2e-08
-100 | 1000V | -99.9999 | 6 | [5 empty] | -0.72648 | 5e-08
-200 | 1000V | -199.9997 | 5.5 | [5 empty] | 0 | 3e-08
-500 | 1000V | -500.0012 | 8.2 | [5 empty] | 0 | 2e-08
-1000 | 1000V | -1000.0024 | 7.6 |  |  |  | 2 |  | -0.066667 | 2e-08
100 | 1000V | 100 | 6 | [5 empty] | 1 | 6e-08
200 | 1000V | 200.0001 | 5.5 | [5 empty] | 0.25 | 3e-08
500 | 1000V | 500.0013 | 8.2 | [5 empty] | 0 | 2e-08
1000 | 1000V | 1000.0027 | 7.6 |  |  |  | 2 |  | 0.052704 | 1e-08
OHM PERFORMANCE TEST
OHM Test 4W | Calibrator ohm | Measured ohm | Source Unc. ppm | Lower Limit ohm | Upper Limit ohm | Measured ppm | DUT Transfer STB ppm | Test Result % of SPEC | MEAS SDEV ppm
1 Ω | 0.99962 | 0.9996 | 85 |  |  |  | 20 |  | 3.9315 | 1.65e-06
1.9 Ω | 1.9 | 1.8999 | 85 |  |  |  | 15 |  | 1.4441 | 9e-07
10 Ω | 10.0003 | 10.0003 | 26 |  |  |  | 5 |  | 0.18782 | 1.7e-07
19 Ω | 19.0037 | 19.0038 | 24 |  |  |  | 5 |  | 0.11093 | 8e-08
100 Ω | 100.0049 | 100.0052 | 15 |  |  |  | 3 |  | 0.033332 | 3e-08
190 Ω | 190.0021 | 190.0031 | 15 |  |  |  | 3 |  | 0.027739 | 2e-08
1 kΩ | 999.9422 | 999.9445 | 11 |  |  |  | 3 |  | 0.052708 | 3e-08
1.9 kΩ | 1899.9031 | 1899.9045 | 11 |  |  |  | 3 |  | 0.017545 | 2e-08
10 kΩ | 10000.399 | 10000.415 | 9 |  |  |  | 3 |  | 0.06009 | 3e-08
19 kΩ | 18999.631 | 18999.637 | 9 |  |  |  | 3 |  | 0.038237 | 2e-08
100 kΩ | 99998.19 | 99998.25 | 11 |  |  |  | 5 |  | 0.070712 | 2e-08
190 kΩ | 189993.6 | 189993.64 | 11 |  |  |  | 5 |  | 0.073394 | 2e-08
1 MΩ | 999904.9 | 999905.3 | 16 |  |  |  | 8 |  | 0.2261 | 5e-08
1.9 MΩ | 1899846.3 | 1899832.1 | 17 |  |  |  | 8 |  | 0.045584 | 4e-08
10MΩ | 9998563 | 9998858 | 33 |  |  |  | 12 |  | 0.1323 | 8e-08
19 MΩ | 18999517 | 19000280 | 43 |  |  |  | 12 |  | 0.20553 | 9e-08
100MΩ(2W) | 100007650 | 100019780 | 110 |  |  |  | 180 |  | 3.0797 | 3.8e-07
OHM Zero 4W | DUT Range | Measured ohm | Source Unc. ohm | Lower Limit ohm | Upper Limit ohm | Measured | DUT Transfer STB ppm | Test Result % of SPEC | MEAS SDEV ohm
0 Ω | 10 Ω | -0.000102 | 5e-05 |  |  | N/A |  |  | 1.1956e-05 | 1.629e-06
0 Ω | 100 Ω | 9e-05 | 5e-05 |  |  | N/A |  |  | 5e-06 | 2.82e-06
0 Ω | 1 KΩ | 0.001 | 5e-05 |  |  | N/A |  |  | 4.4096e-05 | 2.67e-05
0 Ω | 10 KΩ | 0.01 | 5e-05 |  |  | N/A |  |  | 0.0005 | 0.000236
0 Ω | 100 KΩ | 0.07 | 5e-05 |  |  | N/A |  |  | 0.0033333 | 0.000801
0 Ω | 1 MΩ | 0.7 | 5e-05 |  |  | N/A |  |  | 0.033333 | 0.00617
0 Ω | 10 MΩ | 0 | 5e-05 |  |  | N/A |  |  | 0.52705 | 0.0522
0 Ω | 100 MΩ(2W) | -10 | 5e-05 |  |  | N/A |  |  | 0 | 0.508
OHM Zero 2W | DUT Range | Measured ohm | Source Unc. ohm | Lower Limit ohm | Upper Limit ohm | Measured | DUT Transfer STB ppm | Test Result % of SPEC | MEAS SDEV ohm
0 Ω | 10 Ω | 0.45068 | 5e-05 |  |  | N/A |  |  | 0.00016713 | 1.987e-06
0 Ω | 100 Ω | 0.44681 | 5e-05 |  |  | N/A |  |  | 6.4614e-05 | 2.7e-06
0 Ω | 1 KΩ | 0.4476 | 5e-05 |  |  | N/A |  |  | 6.6667e-05 | 2.23e-05
0 Ω | 10 KΩ | 0.458 | 5e-05 |  |  | N/A |  |  | 0.00033333 | 0.000289
0 Ω | 100 KΩ | 0.52 | 5e-05 |  |  | N/A |  |  | 0.0052705 | 0.00064
0 Ω | 1 MΩ | 1.1 | 5e-05 |  |  | N/A |  |  | 0 | 0.0064
0 Ω | 10 MΩ | 1 | 5e-05 |  |  | N/A |  |  | 0 | 0.0526
0 Ω | 100 MΩ | -10 | 5e-05 |  |  | N/A |  |  | 0 | 0.603
ACV PERFORMANCE TEST
ACV Linearity Checks | DUT Range | Measured Vac | Source Unc. ppm | Lower Limit Vac | Upper Limit Vac | Measured ppm of Input | DUT Transfer STB ppm | Test Result % of SPEC | MEAS SDEV ppm
1VAC/1kHz | 10 VAC | 0.9989 | 77 |  |  |  | 10 |  | 82.5698 | 4.56e-06
2VAC/1kHz | 10 VAC | 1.9994 | 73.5 |  |  |  | 10 |  | 8.7029 | 1.64e-06
5VAC/1kHz | 10 VAC | 4.9998 | 84 |  |  |  | 10 |  | 1.3334 | 1.14e-06
10VAC/1kHz | 10 VAC | 10 | 77 |  |  |  | 10 |  | 1.453 | 1.12e-06
12VAC/1kHz | 10 VAC | 12.0001 | 75.833 |  |  |  | 10 |  | 0.80985 | 1.48e-06
15VAC/1kHz | 10 VAC | 15.0002 | 74.667 |  |  |  | 10 |  | 0.92961 | 1.03e-06
19VAC/1kHz | 10 VAC | 19.0003 | 73.684 |  |  |  | 10 |  | 0.87718 | 1.14e-06
ACV Performance Tset |  | Measured Vac | Source Unc. ppm | Lower Limit Vac | Upper Limit Vac | Measured ppm | DUT Transfer STB ppm | Test Result % of SPEC | MEAS SDEV ppm
0.001 V, 10 Hz |  | 0.00099864 | 5500 |  |  |  | 20 |  | 116.1321 | 3.44e-05
0.001 V, 20 Hz |  | 0.00099885 | 5200 |  |  |  | 20 |  | 110.2404 | 4e-05
0.001 V, 30 Hz |  | 0.00099918 | 5200 |  |  |  | 20 |  | 576.7758 | 8.47e-06
0.001 V, 40 Hz |  | 0.00099765 | 5100 |  |  |  | 20 |  | 46.866 | 3.74e-05
0.001 V, 55 Hz |  | 0.00099756 | 5100 |  |  |  | 20 |  | 52.9127 | 2.32e-05
0.001 V, 300 Hz |  | 0.00099763 | 5100 |  |  |  | 20 |  | 122.845 | 0.000227
0.001 V, 1 KHz |  | 0.00099754 | 5100 |  |  |  | 20 |  | 37.9895 | 4.12e-05
0.001 V, 10 KHz |  | 0.0010115 | 5100 |  |  |  | 20 |  | 151605.387 | 0.01018
0.001 V, 20 KHz |  | 0.00118 | 5100 |  |  |  | 20 |  | 41991.5803 | 0.01962
0.001 V, 30 KHz |  | 0.0010032 | 5340 |  |  |  | 20 |  | 199717.6677 | 0.0253
0.001 V, 50 KHz |  | 0.00099997 | 5340 |  |  |  | 20 |  | 81304.5663 | 0.000787
0.001 V, 100 KHz |  | 0.0010213 | 8800 |  |  |  | 50 |  | 8833.7893 | 4.73e-05
0.001 V, 200 KHz |  | 0.00099828 | 16100 |  |  |  | 100 |  | 108.597 | 4.46e-05
0.001 V, 300 KHz |  | 0.00099519 | 16100 |  |  |  | 100 |  | 53.2762 | 2.59e-05
0.001 V, 500 KHz |  | 0.0010014 | 31500 |  |  |  | 200 |  | 62.3201 | 4.33e-05
0.001 V, 1 MHz |  | 0.00099523 | 33000 |  |  |  | 300 |  | 184.5542 | 3.66e-05
0.01 V, 10 Hz |  | 0.0099993 | 1100 |  |  |  | 20 |  | 1.3945 | 2.47e-06
0.01 V, 20 Hz |  | 0.0099985 | 800 |  |  |  | 20 |  | 1.7403 | 1.91e-06
0.01 V, 30 Hz |  | 0.0099983 | 800 |  |  |  | 20 |  | 3.8231 | 1.01e-06
0.01 V, 40 Hz |  | 0.009998 | 700 |  |  |  | 20 |  | 0.97202 | 1.51e-06
0.01 V, 55 Hz |  | 0.0099981 | 700 |  |  |  | 20 |  | 1.8712 | 2.98e-06
0.01 V, 300 Hz |  | 0.0099981 | 700 |  |  |  | 20 |  | 2.45 | 4.02e-06
0.01 V, 1 KHz |  | 0.0099981 | 700 |  |  |  | 20 |  | 1.9225 | 3.32e-06
0.01 V, 10 KHz |  | 0.0099992 | 700 |  |  |  | 20 |  | 0.8334 | 8.6e-07
0.01 V, 20 KHz |  | 0.01 | 700 |  |  |  | 20 |  | 10.7792 | 1.82e-05
0.01 V, 30 KHz |  | 0.010001 | 940 |  |  |  | 20 |  | 1.6582 | 2.08e-06
0.01 V, 50 KHz |  | 0.010002 | 940 |  |  |  | 20 |  | 1.2691 | 8.5e-07
0.01 V, 100 KHz |  | 0.010006 | 1600 |  |  |  | 50 |  | 1.6904 | 8.7e-07
0.01 V, 200 KHz |  | 0.0099988 | 2600 |  |  |  | 100 |  | 2.7389 | 3.46e-06
0.01 V, 300 KHz |  | 0.0099961 | 2600 |  |  |  | 100 |  | 2.9498 | 3.66e-06
0.01 V, 500 KHz |  | 0.010006 | 4500 |  |  |  | 200 |  | 2.1201 | 2.65e-06
0.01 V, 1 MHz |  | 0.010043 | 6000 |  |  |  | 300 |  | 2.1122 | 1.67e-06
0.1 V, 10 Hz |  | 0.10002 | 660 |  |  |  | 20 |  | 7.1403 | 7.68e-06
0.1 V, 20 Hz |  | 0.1 | 300 |  |  |  | 20 |  | 1.74 | 2.16e-06
0.1 V, 30 Hz |  | 0.1 | 300 |  |  |  | 20 |  | 1.1666 | 2.94e-06
0.1 V, 40 Hz |  | 0.1 | 195 |  |  |  | 20 |  | 1.6159 | 2.29e-06
0.1 V, 55 Hz |  | 0.1 | 195 |  |  |  | 20 |  | 2.1081 | 2.27e-06
0.1 V, 300 Hz |  | 0.1 | 195 |  |  |  | 20 |  | 1.2018 | 1.56e-06
0.1 V, 1 KHz |  | 0.1 | 195 |  |  |  | 20 |  | 1.1666 | 1.14e-06
0.1 V, 10 KHz |  | 0.1 | 195 |  |  |  | 20 |  | 0.7071 | 3.1e-07
0.1 V, 20 KHz |  | 0.1 | 195 |  |  |  | 20 |  | 1.5092 | 2.2e-06
0.1 V, 30 KHz |  | 0.1 | 400 |  |  |  | 20 |  | 0.52704 | 4.9e-07
0.1 V, 50 KHz |  | 0.1 | 400 |  |  |  | 20 |  | 1.3017 | 5.3e-07
0.1 V, 100 KHz |  | 0.1 | 1050 |  |  |  | 50 |  | 1.4813 | 6.1e-07
0.1 V, 200 KHz |  | 0.10002 | 1240 |  |  |  | 100 |  | 0.72634 | 5.7e-07
0.1 V, 300 KHz |  | 0.099996 | 1240 |  |  |  | 100 |  | 0.60095 | 7.6e-07
0.1 V, 500 KHz |  | 0.099964 | 1900 |  |  |  | 200 |  | 0.44112 | 6.4e-07
0.1 V, 1 MHz |  | 0.10042 | 4000 |  |  |  | 300 |  | 1.3278 | 7e-07
1 V, 10 Hz |  | 1.0001 | 600 |  |  |  | 10 |  | 9.9459 | 4.56e-06
1 V, 20 Hz |  | 1 | 180 |  |  |  | 10 |  | 1.9649 | 2.86e-06
1 V, 30 Hz |  | 1 | 180 |  |  |  | 10 |  | 1.0541 | 2.63e-06
1 V, 40 Hz |  | 0.99999 | 77 |  |  |  | 10 |  | 0.70712 | 2.53e-06
1 V, 55 Hz |  | 0.99998 | 77 |  |  |  | 10 |  | 3.7007 | 2.01e-06
1 V, 300 Hz |  | 0.99999 | 77 |  |  |  | 10 |  | 1.2361 | 1.85e-06
1 V, 1 KHz |  | 1 | 77 |  |  |  | 10 |  | 0.86602 | 9.2e-07
1 V, 10 KHz |  | 0.99999 | 77 |  |  |  | 10 |  | 0.70712 | 1.01e-06
1 V, 20 KHz |  | 1 | 77 |  |  |  | 10 |  | 0.8819 | 6.8e-07
1 V, 30 KHz |  | 1 | 140 |  |  |  | 10 |  | 0.866 | 5.2e-07
1 V, 50 KHz |  | 1.0001 | 140 |  |  |  | 20 |  | 1.2018 | 5.8e-07
1 V, 100 KHz |  | 0.99984 | 310 |  |  |  | 30 |  | 1.2249 | 8.8e-07
1 V, 200 KHz |  | 0.99952 | 550 |  |  |  | 70 |  | 0.72683 | 5.8e-07
1 V, 300 KHz |  | 1 | 550 |  |  |  | 70 |  | 1.0138 | 6.4e-07
1 V, 500 KHz |  | 1 | 1400 |  |  |  | 100 |  | 0.52704 | 6.6e-07
1 V, 1 MHz |  | 1 | 3000 |  |  |  | 200 |  | 0.97179 | 1.45e-06
10 V, 10 Hz |  | 10.0011 | 600 |  |  |  | 10 |  | 14.0172 | 6.95e-06
10 V, 20 Hz |  | 10.0001 | 180 |  |  |  | 10 |  | 2.6352 | 2.67e-06
10 V, 30 Hz |  | 10 | 180 |  |  |  | 10 |  | 0.70711 | 2.81e-06
10 V, 40 Hz |  | 9.9999 | 77 |  |  |  | 10 |  | 1.9003 | 2.53e-06
10 V, 55 Hz |  | 9.9999 | 77 |  |  |  | 10 |  | 4.7959 | 1e-06
10 V, 300 Hz |  | 9.9999 | 77 |  |  |  | 10 |  | 2.5495 | 1.7e-06
10 V, 1 KHz |  | 10 | 77 |  |  |  | 10 |  | 1 | 8.9e-07
10 V, 10 KHz |  | 9.9998 | 77 |  |  |  | 10 |  | 1.0541 | 3e-07
10 V, 20 KHz |  | 10.0001 | 77 |  |  |  | 10 |  | 0.7071 | 7.3e-07
10 V, 30 KHz |  | 10.0002 | 140 |  |  |  | 10 |  | 0.49999 | 7.1e-07
10 V, 50 KHz |  | 10.0004 | 140 |  |  |  | 20 |  | 0.70708 | 8.2e-07
10 V, 100 KHz |  | 9.9985 | 270 |  |  |  | 30 |  | 1.1306 | 5.9e-07
10 V, 200 KHz |  | 9.9948 | 670 |  |  |  | 70 |  | 0.70747 | 3.3e-07
10 V, 300 KHz |  | 10 | 670 |  |  |  | 70 |  | 0.78174 | 4.7e-07
10 V, 500 KHz |  | 10 | 1700 |  |  |  | 100 |  | 1.0541 | 4e-07
10 V, 1 MHz |  | 10.0001 | 3500 |  |  |  | 200 |  | 2.236 | 2.98e-06
19 V, 10 Hz |  | 19.0068 | 552.63 |  |  |  | 10 |  | 5.2349 | 6e-06
19 V, 20 Hz |  | 19.0018 | 165.79 |  |  |  | 10 |  | 1.4278 | 1.92e-06
19 V, 30 Hz |  | 19.0009 | 165.79 |  |  |  | 10 |  | 1.0229 | 2.56e-06
19 V, 40 Hz |  | 19.0005 | 73.684 |  |  |  | 10 |  | 1.4834 | 2.75e-06
19 V, 55 Hz |  | 19.0002 | 73.684 |  |  |  | 10 |  | 3.5284 | 1.41e-06
19 V, 300 Hz |  | 19 | 73.684 |  |  |  | 10 |  | 1.6246 | 2.09e-06
19 V, 1 KHz |  | 19.0001 | 73.684 |  |  |  | 10 |  | 1.2281 | 1.11e-06
19 V, 10 KHz |  | 19 | 73.684 |  |  |  | 10 |  | 0.51149 | 6.7e-07
19 V, 20 KHz |  | 19.0005 | 73.684 |  |  |  | 10 |  | 0.68509 | 6.3e-07
19 V, 30 KHz |  | 19.0007 | 130.53 |  |  |  | 10 |  | 0.5263 | 6.7e-07
19 V, 50 KHz |  | 19.0012 | 130.53 |  |  |  | 20 |  | 0.55475 | 3.9e-07
19 V, 100 KHz |  | 18.9977 | 251.05 |  |  |  | 30 |  | 0.51155 | 5.1e-07
19 V, 200 KHz |  | 18.9911 | 589.47 |  |  |  | 70 |  | 0.58871 | 8.1e-07
19 V, 300 KHz |  | 19.001 | 589.47 |  |  |  | 70 |  | 1.4198 | 8.1e-07
19 V, 500 KHz |  | 19.0013 | 1463.2 |  |  |  | 100 |  | 0.79428 | 7.1e-07
19 V, 1 MHz |  | 19.0034 | 3073.7 |  |  |  | 200 |  | 12.2537 | 1.27e-06
100 V, 10 Hz |  | 100.0116 | 600 |  |  |  | 10 |  | 7.896 | 1.054e-05
100 V, 20 Hz |  | 100.0015 | 180 |  |  |  | 10 |  | 3.1666 | 2.24e-06
100 V, 30 Hz |  | 99.9998 | 180 |  |  |  | 10 |  | 1.118 | 2.21e-06
100 V, 40 Hz |  | 99.9986 | 85 |  |  |  | 10 |  | 1.3945 | 2.78e-06
100 V, 55 Hz |  | 99.9987 | 85 |  |  |  | 10 |  | 2.7285 | 1.12e-06
100 V, 300 Hz |  | 99.9981 | 85 |  |  |  | 10 |  | 1.8709 | 1.93e-06
100 V, 1 KHz |  | 99.9977 | 85 |  |  |  | 10 |  | 1.5 | 1.09e-06
100 V, 10 KHz |  | 99.9979 | 85 |  |  |  | 10 |  | 0.52706 | 5.1e-07
100 V, 20 KHz |  | 99.9973 | 85 |  |  |  | 10 |  | 0.88194 | 5.4e-07
100 V, 30 KHz |  | 99.9983 | 240 |  |  |  | 10 |  | 0.44097 | 5.6e-07
100 V, 50 KHz |  | 99.9978 | 240 |  |  |  | 20 |  | 0.86604 | 1.22e-06
100 V, 100 KHz |  | 99.9953 | 600 |  |  |  | 30 |  | 1.9862 | 9.5e-07
100 V, 200 KHz |  | 100.0012 | 2600 |  |  |  | 50 |  | 5.1988 | 5.08e-06
1000 V, 55 Hz |  | 1000.02 | 79 |  |  |  | 15 |  | 2.8284 | 1.52e-06
1000 V, 300 Hz |  | 1000.026 | 79 |  |  |  | 15 |  | 0.66665 | 2.12e-06
700 V, 1 KHz |  | 699.997 | 80.714 |  |  |  | 50 |  | 0.75293 | 1.83e-06
DCI PERFORMANCE TEST
DCI Test | Expected Value | Measured A | Source Unc. ppm | Lower Limit A | Upper Limit A | Measured ppm | DUT Transfer STB ppm | Test Result % of SPEC | MEAS SDEV ppm
1 µADC | 1e-06 | 9.972e-07 | 10073.205 | [5 empty] | 50.1404 | 1.25e-05
10 µADC | 1e-05 | 9.9976e-06 | 1045.2509 | [5 empty] | 5.0012 | 1.04e-06
50 µADC | 5e-05 | 4.9999e-05 | 245 | [5 empty] | 1 | 2.7e-07
100 µADC | 0.0001 | 0.0001 | 145 |  |  |  | 7 |  | 0 | 9e-08
-100 µADC | -0.0001 | -0.00010001 | 145 |  |  |  | 7 |  | 0 | 1e-07
-50 µADC | -5e-05 | -5.0005e-05 | 245 | [5 empty] | 0 | 2e-07
-10 µADC | -1e-05 | -1.0003e-05 | -1044.6657 | [5 empty] | 0 | 9e-07
0.1 mADC | 0.0001 | 9.9976e-05 | 145 | [5 empty] | 0 | 1.31e-06
0.5 mADC | 0.0005 | 0.00049999 | 65 | [5 empty] | 0 | 2.1e-07
1 mADC | 0.001 | 0.001 | 55 |  |  |  | 7 |  | 0 | 1.1e-07
-1 mADC | -0.001 | -0.001 | 55 |  |  |  | 7 |  | -0.49998 | 1.3e-07
-0.5 mADC | -0.0005 | -0.00050004 | 65 | [5 empty] | -0.99993 | 2.4e-07
1 mADC | 0.001 | 0.00099999 | 55 | [5 empty] | 5.0001 | 1.11e-06
5 mADC | 0.005 | 0.005 | 65 | [5 empty] | 1 | 2.3e-07
10 mADC | 0.01 | 0.01 | 55 |  |  |  | 7 |  | 0 | 1.2e-07
-10 mADC | -0.01 | -0.01 | 55 |  |  |  | 7 |  | -0.57735 | 9e-08
-5 mADC | -0.005 | -0.005 | 65 | [5 empty] | 0 | 3.5e-07
10 mADC | 0.01 | 0.0099997 | 55 | [5 empty] | 5.0002 | 1.32e-06
50 mADC | 0.05 | 0.05 | 75 | [5 empty] | 1 | 2.5e-07
100 mADC | 0.1 | 0.1 | 65 |  |  |  | 7 |  | 0.5 | 1.1e-07
-100 mADC | -0.1 | -0.1 | 65 |  |  |  | 7 |  | -0.5 | 1.4e-07
-50 mADC | -0.05 | -0.05 | 75 | [5 empty] | -1 | 2e-07
0.1 ADC | 0.1 | 0.1 | 65 | [5 empty] | 5.7734 | 1.13e-06
0.5 ADC | 0.5 | 0.5 | 135 | [5 empty] | 1.1547 | 2e-07
-0.5 ADC | -0.5 | -0.5 | 135 | [5 empty] | -0.99999 | 2e-07
-1.0 ADC | -1 | -0.99998 | 105 |  |  |  | 20 |  | -0.50001 | 1.6e-07
1.0 ADC | 1 | 0.99998 | 105 |  |  |  | 20 |  | 0 | 1.3e-07
ACI PERFORMANCE TEST
ACI Test | Expected Value | Measured Aac | Source Unc. ppm | Lower Limit Aac | Upper Limit Aac | Measured ppm | DUT Transfer STB ppm | Test Result % of SPEC | MEAS SDEV ppm
10 µA, 10 Hz | 1e-05 | 1.0004e-05 | 3648.5953 | [5 empty] | 9.5706 | 1.012e-05
10 µA, 20 Hz | 1e-05 | 1.0003e-05 | 2849.0171 | [5 empty] | 5.7715 | 3.9e-06
10 µA, 30 Hz | 1e-05 | 1.0006e-05 | 2848.2483 | [5 empty] | 120.2035 | 4.19e-06
10 µA, 40 Hz | 1e-05 | 1.0006e-05 | 2118.6652 | [5 empty] | 4.9969 | 5.39e-06
10 µA 55 Hz | 1e-05 | 1.0006e-05 | 2118.7288 | [5 empty] | 0 | 6.6e-06
10 µA, 300 Hz | 1e-05 | 1.0006e-05 | 2118.6547 |  |  | -40 |  |  | 9.5682 | 8.78e-06
10 µA, 1 KHz | 1e-05 | 1.0006e-05 | 2118.6229 | [5 empty] | 16.3193 | 2.1e-05
10 µA, 5 KHz | 1e-05 | 1.0006e-05 | 5496.9217 | [5 empty] | 8.1604 | 3.87e-06
10 µA, 10 KHz | 1e-05 | 1.0003e-05 | 11496.0339 | [5 empty] | 63.9441 | 2.57e-05
100 µA, 10 Hz | 0.0001 | 0.00010002 | 950 | 9.9939e-05 | 0.00010006 | -40 | 50 |  | 6.8482 | 8.44e-06
100 µA, 20 Hz | 0.0001 | 0.00010001 | 600 |  |  |  | 50 |  | 6.6011 | 1.97e-06
100 µA, 30 Hz | 0.0001 | 0.00010001 | 600 | 9.9939e-05 | 0.00010006 |  | 50 |  | 1.2909 | 2.06e-06
100 µA, 40 Hz | 0.0001 | 0.00010001 | 320 |  |  |  | 50 |  | 0.99991 | 1.57e-06
100 µA 55 Hz | 0.0001 | 0.00010001 | 320 | 9.9939e-05 | 0.00010006 | -40 | 50 |  | 0.57731 | 2.1e-06
100 µA, 300 Hz | 0.0001 | 0.00010001 | 320 |  |  |  | 50 |  | 3.4032 | 1.76e-06
100 µA, 1 KHz | 0.0001 | 0.0001 | 320 | 9.9939e-05 | 0.00010006 |  | 50 |  | 1.7078 | 2.72e-06
100 µA, 5 KHz | 0.0001 | 0.0001 | 0.1 |  |  |  | 100 |  | 0.49999 | 5.6e-07
100 µA, 10 KHz | 0.0001 | 0.00010002 | 0.25 | 9.9939e-05 | 0.00010006 |  | 300 |  | 1.2908 | 5.9e-07
1mA, 10 Hz | 0.001 | 0.0010002 | 700 |  |  |  | 40 |  | 5.6852 | 1.003e-05
1 mA, 20 Hz | 0.001 | 0.0010001 | 390 | 9.9939e-05 | 0.00010006 |  | 40 |  | 3.4998 | 1.19e-06
1 mA, 30 Hz | 0.001 | 0.0010001 | 390 |  |  |  | 40 |  | 0.81645 | 2.5e-06
1 mA, 40 Hz | 0.001 | 0.001 | 160 | 9.9939e-05 | 0.00010006 |  | 40 |  | 1.2583 | 2.15e-06
1 mA, 55 Hz | 0.001 | 0.001 | 160 |  |  |  | 40 |  | 1.2909 | 1.68e-06
1 mA, 300 Hz | 0.001 | 0.001 | 160 | 9.9939e-05 | 0.00010006 |  | 40 |  | 2.1602 | 2.4e-06
1 mA, 1 KHz | 0.001 | 0.001 | 160 |  |  |  | 40 |  | 0.81647 | 1.02e-06
1 mA, 5 KHz | 0.001 | 0.001 | 1000 | 9.9939e-05 | 0.00010006 |  | 70 |  | 0.57734 | 6.7e-07
1 mA, 10 KHz | 0.001 | 0.00099999 | 0.25 |  |  |  | 200 |  | 0.57736 | 6.1e-07
10 mA, 10 Hz | 0.01 | 0.010002 | 700 | 9.9939e-05 | 0.00010006 |  | 40 |  | 1.7075 | 1.043e-05
10 mA, 20 Hz | 0.01 | 0.010001 | 390 |  |  |  | 40 |  | 0.49996 | 1.33e-06
10 mA, 30 Hz | 0.01 | 0.010001 | 390 | 9.9939e-05 | 0.00010006 |  | 40 |  | 0.81645 | 2.23e-06
10 mA, 40 Hz | 0.01 | 0.010001 | 160 |  |  |  | 40 |  | 7.4996 | 1.59e-06
10 mA, 55 Hz | 0.01 | 0.01 | 160 | 9.9939e-05 | 0.00010006 |  | 40 |  | 0.49998 | 1.86e-06
10 mA, 300 Hz | 0.01 | 0.01 | 160 |  |  |  | 40 |  | 0.81646 | 1.78e-06
10 mA, 1 KHz | 0.01 | 0.010001 | 160 | 9.9939e-05 | 0.00010006 |  | 40 |  | 1.1546 | 1.12e-06
10 mA, 5 KHz | 0.01 | 0.010001 | 1000 |  |  |  | 70 |  | 0 | 8.3e-07
10 mA, 10 KHz | 0.01 | 0.0099999 | 0.25 | 9.9939e-05 | 0.00010006 |  | 200 |  | 1.291 | 3.4e-07
100 mA, 10 Hz | 0.1 | 0.10002 | 700 |  |  |  | 40 |  | 6.2903 | 6.96e-06
100 mA, 20 Hz | 0.1 | 0.10001 | 390 | 9.9939e-05 | 0.00010006 |  | 40 |  | 1.4141 | 1.8e-06
100 mA, 30 Hz | 0.1 | 0.10001 | 390 |  |  |  | 40 |  | 1.2582 | 1.74e-06
100 mA, 40 Hz | 0.1 | 0.10001 | 160 | 9.9939e-05 | 0.00010006 |  | 40 |  | 3.0954 | 1.92e-06
100 mA, 55 Hz | 0.1 | 0.10001 | 160 |  |  |  | 40 |  | 1.8928 | 2.16e-06
100 mA, 300 Hz | 0.1 | 0.10001 | 160 | 9.9939e-05 | 0.00010006 |  | 40 |  | 0.95735 | 1.99e-06
100 mA, 1 KHz | 0.1 | 0.10001 | 160 |  |  |  | 40 |  | 0.49995 | 1e-06
100 mA, 5 KHz | 0.1 | 0.099999 | 0.1 | 9.9939e-05 | 0.00010006 |  | 70 |  | 1.291 | 5.2e-07
100 mA, 10 KHz | 0.1 | 0.099999 | 0.25 |  |  |  | 200 |  | 1.893 | 6.6e-07
1 A, 10 Hz | 1 | 0.99996 | -1 | 9.9939e-05 | 0.00010006 |  | 40 |  | 12.5537 | 1.013e-05
1 A, 20 Hz | 1 | 0.99998 | -1 |  |  |  | 40 |  | 6.0554 | 1.37e-06
1 A, 30 Hz | 1 | 0.99997 | -1 | 9.9939e-05 | 0.00010006 |  | 40 |  | 0.57737 | 1.3e-06
1 A, 40 Hz | 1 | 1.0001 | 640 |  |  |  | 40 |  | 5.0658 | 1.1e-06
1 A, 55 Hz | 1 | 1.0001 | 640 | 9.9939e-05 | 0.00010006 |  | 40 |  | 1.2909 | 1.6e-06
1 A, 300 Hz | 1 | 1.0001 | 640 |  |  |  | 40 |  | 3.1089 | 2.11e-06
1 A, 1 KHz | 1 | 1.0001 | 640 | 9.9939e-05 | 0.00010006 |  | 40 |  | 0.95738 | 1.28e-06
1 A, 5 KHz | 1 | 0.99988 | 800 |  |  |  | 70 |  | 0 | 4.9e-07
1 A, 10 KHz | 1 | 1.0008 | 0.82 | 9.9939e-05 | 0.00010006 |  | 200 |  | 1.1538 | 2.7e-07
Test Summary | Initialization | DCV | OHM | ACV | DCI | ACI |  | Total
Time escape minute | 0.19 | 206.9 | 249.15 | 683.28 | 40.4 | 186.12