## Sheet
PERFORMANCE TEST
DUT
Manufacturer |  | Datron Instruments |  |  | S/N |  | 31255
Model Number |  | 4950 |  |  | Test date |  | 23-09-23 14:36:22
Meter Info |  | Datron 4950 STD |  |  | Ambient Temp. |  | 23.0
Options |  | 1,1,1 |  |  | CERT_AMB? |  | 23
Firmware |  | 890197-02.07 |  |  | TEMP? |  | 42.22
Date of Baseline Cal |  | "1999-SEP-16" |  |  | TEMP? BASE |  | 2.22
Date of Certified Cal |  | "2023-SEP-07" |  |  | TEMP? CERTIFIED |  | -0.74
CNFTST? |  | 0 |  |  | STANDBY? |  | "00:00"
DUT Confidence level |  | 30 Days Transfer Satbility
Calibrator
Reference standard |  | Options |  | Unc | Calibration Date |  | Due Date
Fluke 5700A |  |  |  | -- | -- |  | --
ADRmu12 |  | 10.000000 VDC |  | -- | -- |  | --
Fluke 742A-1 8927002 |  | 1.0000195 ohm |  | -- | -- |  | --
Fluke 742A-10k 7815006 |  | 10.000133 kohm
MFC
MFC Calibrate Date |  | 030622 |  |  | MFC last calibrated |  | 477 days
Calibrate Date WBFLAT |  | 100589 |  |  | MFC since WBFLAT |  | 0 days
Calibrate Date WB Gain |  | 250190 |  |  | MFC since WBGAIN |  | 0 days
MFC Calibrate Date Zero |  | 270823 |  |  | MFC since DCV Zero |  | 27 days
MFC Confidence level |  | 5700A 99% 90 Days |  |  | CAL CONST 11V Range +zero |  | 398.1791
CAL CONST 6.5V |  | +6.529921596265443E+00 |  |  | CAL CONST 11V Range -zero |  | 398.1783
CAL CONST 13V reference voltage |  | +1.306136884003510E+01 |  |  | CAL CONST, Zero calibration temperature |  | 23
CAL CONST 10KOHM true output resistance |  | +1.000039902274076E+04 |  |  | CAL CONST, ALL calibration temp |  | 23
CAL SONST 10kOHM standard resistance |  | +1.000059914344082E+04 |  |  | CAL_CONST? CHECK, WB_FLAT_TEMP |  | 23.9
CAL CONST 1 OHM |  | +9.996206897944441E-01 |  |  | CAL CONST,WBAND_TEMP
ETIME? |  | 6051 days 1.70 hours |  |  | CAL TEMP? CHECK
ONTIME? |  | 0 days 4.90 hours |  |  | CAL CONST,DACLIN |  | 0.359
DCV PERFORMANCE TEST
DCV Test VDC | Expected Value | Measured μV | Source Unc. μV | Lower Limit μV | Upper Limit μV | Deviation | DUT Transfer STB ppm | Test Result % of SPEC | MEAS SDEV μV
0.0mVDC | 0 | -0.01 | 7.5e-07 |  |  | N/A |  |  | 0.029345 | 0.00592
0.0VDC | 0 | 0.9 | 7.5e-07 |  |  | N/A |  |  | 0.05 | 0.0134
0.0VDC | 0 | 6 | 7.5e-07 |  |  | N/A |  |  | 0 | 0.068
0.0VDC | 0 | 70 | 7.5e-07 |  |  | N/A |  |  | 0 | 0.723
0.0VDC | 0 | 400 | 7.5e-07 |  |  | N/A |  |  | 44.0959 | 4.39
DCV Test VDC | DMM Range | Measured V | Source Unc. ppm | Lower Limit V | Upper Limit V | Measured ppm | DUT Transfer STB ppm | Test Result % of SPEC | MEAS SDEV ppm
0.01 | 100mV | 0.01 | 81.5 | [5 empty] | 1.0138 | 5.6e-07
0.02 | 100mV | 0.02 | 44 | [5 empty] | 0.8819 | 2.7e-07
0.05 | 100mV | 0.05 | 21.5 | [5 empty] | 0.16667 | 1.3e-07
0.1 | 100mV | 0.1 | 14 |  |  |  | 3 |  | 0.22791 | 6e-08
-0.01 | 100mV | -0.0099995 | -81.5 | [5 empty] | -1.3229 | 6.1e-07
-0.02 | 100mV | -0.02 | -44 | [5 empty] | -0.30047 | 3.2e-07
-0.05 | 100mV | -0.049999 | 21.5 | [5 empty] | -0.31798 | 1.3e-07
-0.1 | 100mV | -0.099999 | 14 |  |  |  | 3 |  | -0.070711 | 6e-08
0.1 | 1V | 0.1 | 14 | [5 empty] | 1.0929 | 1.5e-07
0.2 | 1V | 0.2 | 10.25 | [5 empty] | 0.5 | 7e-08
0.5 | 1V | 0.5 | 5.9 | [5 empty] | 0.22608 | 4e-08
1 | 1V | 1 | 4.7 |  |  |  | 1.5 |  | 0.086602 | 3e-08
-0.1 | 1V | -0.099998 | 14 | [5 empty] | -1.2361 | 1.1e-07
-0.2 | 1V | -0.2 | 10.25 | [5 empty] | -0.22048 | 6e-08
-0.5 | 1V | -0.5 | 5.9 | [5 empty] | -0.12019 | 3e-08
-1 | 1V | -1 | 4.7 |  |  |  | 1.5 |  | -0.060093 | 2e-08
1 | 10V | 1 | 4.7 | [5 empty] | 0.5 | 8e-08
2 | 10V | 2 | 4.1 | [5 empty] | 0.30046 | 5e-08
5 | 10V | 5 | 4.1 | [5 empty] | 0.14142 | 2e-08
10 | 10V | 10 | 3.8 |  |  |  | 1.5 |  | 0.066667 | 1e-08
19 | 10V | 19 | 3.8158 |  |  |  | 1.5 |  | 0.094883 | 1e-08
-1 | 10V | -0.99999 | 4.7 | [5 empty] | -1.7321 | 6e-08
-2 | 10V | -2 | 4.1 | [5 empty] | -0.25 | 4e-08
-5 | 10V | -5 | 4.1 | [5 empty] | -0.13333 | 2e-08
-10 | 10V | -10 | 3.8 |  |  |  | 0.5 |  | -0.066667 | 2e-08
-19 | 10V | -19 | 3.8158 |  |  |  | 1.5 |  | -0.053357 | 1e-08
10 | 100V | 10.0001 | 3.8 | [5 empty] | 1.4142 | 1e-07
20 | 100V | 20.0003 | 3.8 | [5 empty] | 0.33333 | 5e-08
50 | 100V | 50.0005 | 7 | [5 empty] | 0.10541 | 3e-08
100 | 100V | 100.0009 | 6 |  |  |  | 2 |  | 0.07071 | 2e-08
-10 | 100V | -9.9999 | 3.8 | [5 empty] | -0.88193 | 8e-08
-20 | 100V | -19.9998 | 3.8 | [5 empty] | -0.26353 | 4e-08
-50 | 100V | -49.9995 | 7 | [5 empty] | -0.066667 | 3e-08
-100 | 100V | -99.9992 | 6 |  |  |  | 2 |  | -0.072649 | 2e-08
-100 | 1000V | -99.9998 | 6 | [5 empty] | -0.33333 | 4e-08
-200 | 1000V | -199.9997 | 5.5 | [5 empty] | 0 | 3e-08
-500 | 1000V | -500.0004 | 8.2 | [5 empty] | 0 | 2e-08
-1000 | 1000V | -1000.0011 | 7.6 |  |  |  | 2 |  | -0.092796 | 2e-08
100 | 1000V | 100.0003 | 6 | [5 empty] | 0.72648 | 7e-08
200 | 1000V | 200.0003 | 5.5 | [5 empty] | 0 | 3e-08
500 | 1000V | 500.0009 | 8.2 | [5 empty] | 0 | 3e-08
1000 | 1000V | 1000.0017 | 7.6 |  |  |  | 2 |  | 0.15 | 2e-08
OHM PERFORMANCE TEST
OHM Test 4W | Calibrator ohm | Measured ohm | Source Unc. ppm | Lower Limit ohm | Upper Limit ohm | Measured ppm | DUT Transfer STB ppm | Test Result % of SPEC | MEAS SDEV ppm
1 Ω | 0.99962 | 0.99958 | 85 |  |  |  | 20 |  | 2.0624 | 1.79e-06
1.9 Ω | 1.9 | 1.8999 | 85 |  |  |  | 15 |  | 0.89031 | 8.4e-07
10 Ω | 10.0003 | 10.0003 | 26 |  |  |  | 5 |  | 0.33081 | 1.6e-07
19 Ω | 19.0037 | 19.0038 | 24 |  |  |  | 5 |  | 0.14909 | 8e-08
100 Ω | 100.0049 | 100.0053 | 15 |  |  |  | 3 |  | 0.086598 | 3e-08
190 Ω | 190.0021 | 190.0032 | 15 |  |  |  | 3 |  | 0.05752 | 2e-08
1 kΩ | 999.9422 | 999.9449 | 11 |  |  |  | 3 |  | 0.13018 | 4e-08
1.9 kΩ | 1899.9031 | 1899.9044 | 11 |  |  |  | 3 |  | 0.070179 | 2e-08
10 kΩ | 10000.399 | 10000.411 | 9 |  |  |  | 3 |  | 0.1054 | 4e-08
19 kΩ | 18999.631 | 18999.624 | 9 |  |  |  | 3 |  | 0.074949 | 3e-08
100 kΩ | 99998.19 | 99998.14 | 11 |  |  |  | 5 |  | 0.15899 | 2e-08
190 kΩ | 189993.6 | 189993.46 | 11 |  |  |  | 5 |  | 0.066807 | 2e-08
1 MΩ | 999904.9 | 999898.8 | 16 |  |  |  | 8 |  | 0.32833 | 3e-08
1.9 MΩ | 1899846.3 | 1899822 | 17 |  |  |  | 8 |  | 0.070182 | 4e-08
10MΩ | 9998563 | 9998261 | 33 |  |  |  | 12 |  | 0.15369 | 1.1e-07
19 MΩ | 18999517 | 18999289 | 43 |  |  |  | 12 |  | 0.07797 | 7e-08
100MΩ(2W) | 100007650 | 100004150 | 110 |  |  |  | 180 |  | 1.651 | 2e-07
OHM Zero 4W | DUT Range | Measured ohm | Source Unc. ohm | Lower Limit ohm | Upper Limit ohm | Measured | DUT Transfer STB ppm | Test Result % of SPEC | MEAS SDEV ohm
0 Ω | 10 Ω | -0.000123 | 5e-05 |  |  | N/A |  |  | 1.1035e-05 | 1.61e-06
0 Ω | 100 Ω | 0.00016 | 5e-05 |  |  | N/A |  |  | 0 | 2.33e-06
0 Ω | 1 KΩ | 0.0017 | 5e-05 |  |  | N/A |  |  | 3.3333e-05 | 2.77e-05
0 Ω | 10 KΩ | 0.018 | 5e-05 |  |  | N/A |  |  | 0.00052705 | 0.000261
0 Ω | 100 KΩ | 0.03 | 5e-05 |  |  | N/A |  |  | 0.0060093 | 0.000835
0 Ω | 1 MΩ | 0.4 | 5e-05 |  |  | N/A |  |  | 0.05 | 0.00827
0 Ω | 10 MΩ | 2 | 5e-05 |  |  | N/A |  |  | 0 | 0.0546
0 Ω | 100 MΩ(2W) | 30 | 5e-05 |  |  | N/A |  |  | 0 | 0.559
OHM Zero 2W | DUT Range | Measured ohm | Source Unc. ohm | Lower Limit ohm | Upper Limit ohm | Measured | DUT Transfer STB ppm | Test Result % of SPEC | MEAS SDEV ohm
0 Ω | 10 Ω | 0.30917 | 5e-05 |  |  | N/A |  |  | 0.00018334 | 2.37e-06
0 Ω | 100 Ω | 0.30537 | 5e-05 |  |  | N/A |  |  | 7.0848e-05 | 2.28e-06
0 Ω | 1 KΩ | 0.3068 | 5e-05 |  |  | N/A |  |  | 6.6667e-05 | 3.03e-05
0 Ω | 10 KΩ | 0.324 | 5e-05 |  |  | N/A |  |  | 0.00033333 | 0.000289
0 Ω | 100 KΩ | 0.34 | 5e-05 |  |  | N/A |  |  | 0.0044096 | 0.000789
0 Ω | 1 MΩ | 0.6 | 5e-05 |  |  | N/A |  |  | 0 | 0.00597
0 Ω | 10 MΩ | 2 | 5e-05 |  |  | N/A |  |  | 0.44096 | 0.0521
0 Ω | 100 MΩ | 30 | 5e-05 |  |  | N/A |  |  | 6.0093 | 0.562
ACV PERFORMANCE TEST
ACV Linearity Checks | DUT Range | Measured Vac | Source Unc. ppm | Lower Limit Vac | Upper Limit Vac | Measured ppm of Input | DUT Transfer STB ppm | Test Result % of SPEC | MEAS SDEV ppm
1VAC/1kHz | 10 VAC | 0.99893 | 77 |  |  |  | 10 |  | 81.1732 | 2.5e-06
2VAC/1kHz | 10 VAC | 1.9994 | 73.5 |  |  |  | 10 |  | 10.1409 | 1.57e-06
5VAC/1kHz | 10 VAC | 4.9998 | 84 |  |  |  | 10 |  | 1 | 6.9e-07
10VAC/1kHz | 10 VAC | 10 | 77 |  |  |  | 10 |  | 0.88192 | 1.2e-06
12VAC/1kHz | 10 VAC | 12 | 75.833 |  |  |  | 10 |  | 1.1785 | 9.6e-07
15VAC/1kHz | 10 VAC | 15.0002 | 74.667 |  |  |  | 10 |  | 1.4824 | 1.12e-06
19VAC/1kHz | 10 VAC | 19.0003 | 73.684 |  |  |  | 10 |  | 1.2374 | 1.16e-06
ACV Performance Tset |  | Measured Vac | Source Unc. ppm | Lower Limit Vac | Upper Limit Vac | Measured ppm | DUT Transfer STB ppm | Test Result % of SPEC | MEAS SDEV ppm
0.001 V, 10 Hz |  | 0.0009973 | 5500 |  |  |  | 20 |  | 71.6076 | 5.57e-06
0.001 V, 20 Hz |  | 0.00099747 | 5200 |  |  |  | 20 |  | 18.7557 | 2.51e-06
0.001 V, 30 Hz |  | 0.00099754 | 5200 |  |  |  | 20 |  | 11.3318 | 4.97e-06
0.001 V, 40 Hz |  | 0.00099754 | 5100 |  |  |  | 20 |  | 10.0247 | 4.11e-06
0.001 V, 55 Hz |  | 0.00099785 | 5100 |  |  |  | 20 |  | 7.2805 | 2.06e-06
0.001 V, 300 Hz |  | 0.00099782 | 5100 |  |  |  | 20 |  | 12.2742 | 1.834e-05
0.001 V, 1 KHz |  | 0.00099789 | 5100 |  |  |  | 20 |  | 8.6786 | 1.365e-05
0.001 V, 10 KHz |  | 0.00099793 | 5100 |  |  |  | 20 |  | 7.8336 | 9.33e-06
0.001 V, 20 KHz |  | 0.001008 | 5100 |  |  |  | 20 |  | 24.8021 | 3.33e-05
0.001 V, 30 KHz |  | 0.001008 | 5340 |  |  |  | 20 |  | 4.3745 | 7.1e-06
0.001 V, 50 KHz |  | 0.0010081 | 5340 |  |  |  | 20 |  | 8.591 | 4.26e-06
0.001 V, 100 KHz |  | 0.0010014 | 8800 |  |  |  | 50 |  | 7.0613 | 4.88e-06
0.001 V, 200 KHz |  | 0.00099856 | 16100 |  |  |  | 100 |  | 15.0216 | 8.48e-06
0.001 V, 300 KHz |  | 0.00099559 | 16100 |  |  |  | 100 |  | 8.8582 | 5.34e-06
0.001 V, 500 KHz |  | 0.001002 | 31500 |  |  |  | 200 |  | 4.9902 | 1.401e-05
0.001 V, 1 MHz |  | 0.00099683 | 33000 |  |  |  | 300 |  | 11.7038 | 7.6e-06
0.01 V, 10 Hz |  | 0.0099983 | 1100 |  |  |  | 20 |  | 1.0931 | 2.62e-06
0.01 V, 20 Hz |  | 0.0099973 | 800 |  |  |  | 20 |  | 0.86626 | 1.14e-06
0.01 V, 30 Hz |  | 0.009997 | 800 |  |  |  | 20 |  | 0.50015 | 1.38e-06
0.01 V, 40 Hz |  | 0.0099967 | 700 |  |  |  | 20 |  | 6.4463 | 6.9e-07
0.01 V, 55 Hz |  | 0.0099968 | 700 |  |  |  | 20 |  | 0.8822 | 2.36e-06
0.01 V, 300 Hz |  | 0.0099967 | 700 |  |  |  | 20 |  | 2.2981 | 3e-06
0.01 V, 1 KHz |  | 0.0099968 | 700 |  |  |  | 20 |  | 1.5097 | 1.96e-06
0.01 V, 10 KHz |  | 0.0099979 | 700 |  |  |  | 20 |  | 0.66681 | 7e-07
0.01 V, 20 KHz |  | 0.0099989 | 700 |  |  |  | 20 |  | 2.1797 | 3.14e-06
0.01 V, 30 KHz |  | 0.0099992 | 940 |  |  |  | 20 |  | 0.7818 | 9.2e-07
0.01 V, 50 KHz |  | 0.01 | 940 |  |  |  | 20 |  | 0.97178 | 3.1e-07
0.01 V, 100 KHz |  | 0.010005 | 1600 |  |  |  | 50 |  | 0.8815 | 7.3e-07
0.01 V, 200 KHz |  | 0.0099971 | 2600 |  |  |  | 100 |  | 1.4146 | 8.1e-07
0.01 V, 300 KHz |  | 0.009994 | 2600 |  |  |  | 100 |  | 0.83383 | 1.03e-06
0.01 V, 500 KHz |  | 0.010004 | 4500 |  |  |  | 200 |  | 0.99964 | 1.71e-06
0.01 V, 1 MHz |  | 0.01004 | 6000 |  |  |  | 300 |  | 1.0098 | 1.32e-06
0.1 V, 10 Hz |  | 0.10002 | 660 |  |  |  | 20 |  | 7.8728 | 9.49e-06
0.1 V, 20 Hz |  | 0.1 | 300 |  |  |  | 20 |  | 1.3228 | 1.51e-06
0.1 V, 30 Hz |  | 0.1 | 300 |  |  |  | 20 |  | 1.118 | 3.2e-06
0.1 V, 40 Hz |  | 0.1 | 195 |  |  |  | 20 |  | 0.83332 | 1.83e-06
0.1 V, 55 Hz |  | 0.1 | 195 |  |  |  | 20 |  | 3.1622 | 2.69e-06
0.1 V, 300 Hz |  | 0.1 | 195 |  |  |  | 20 |  | 1.74 | 2.25e-06
0.1 V, 1 KHz |  | 0.1 | 195 |  |  |  | 20 |  | 0.86602 | 1.19e-06
0.1 V, 10 KHz |  | 0.1 | 195 |  |  |  | 20 |  | 1.0929 | 4.8e-07
0.1 V, 20 KHz |  | 0.1 | 195 |  |  |  | 20 |  | 0.66666 | 4.3e-07
0.1 V, 30 KHz |  | 0.1 | 400 |  |  |  | 20 |  | 0.92795 | 3.4e-07
0.1 V, 50 KHz |  | 0.1 | 400 |  |  |  | 20 |  | 0.52705 | 3.8e-07
0.1 V, 100 KHz |  | 0.1 | 1050 |  |  |  | 50 |  | 1.0541 | 7.9e-07
0.1 V, 200 KHz |  | 0.10002 | 1240 |  |  |  | 100 |  | 0.83319 | 5.7e-07
0.1 V, 300 KHz |  | 0.099988 | 1240 |  |  |  | 100 |  | 0.52711 | 7.4e-07
0.1 V, 500 KHz |  | 0.09995 | 1900 |  |  |  | 200 |  | 0.44118 | 8.5e-07
0.1 V, 1 MHz |  | 0.10002 | 4000 |  |  |  | 300 |  | 1.1178 | 6.6e-07
1 V, 10 Hz |  | 1.0001 | 600 |  |  |  | 10 |  | 6.7898 | 8.76e-06
1 V, 20 Hz |  | 1 | 180 |  |  |  | 10 |  | 1.3017 | 1.63e-06
1 V, 30 Hz |  | 1 | 180 |  |  |  | 10 |  | 0.83334 | 3.34e-06
1 V, 40 Hz |  | 0.99999 | 77 |  |  |  | 10 |  | 0.92797 | 1.96e-06
1 V, 55 Hz |  | 0.99999 | 77 |  |  |  | 10 |  | 3.1136 | 2.33e-06
1 V, 300 Hz |  | 0.99999 | 77 |  |  |  | 10 |  | 1.9221 | 2.16e-06
1 V, 1 KHz |  | 1 | 77 |  |  |  | 10 |  | 1.3333 | 1.06e-06
1 V, 10 KHz |  | 0.99999 | 77 |  |  |  | 10 |  | 0.72649 | 4.8e-07
1 V, 20 KHz |  | 1 | 77 |  |  |  | 10 |  | 0.7071 | 3.6e-07
1 V, 30 KHz |  | 1 | 140 |  |  |  | 10 |  | 0.52703 | 5.5e-07
1 V, 50 KHz |  | 1 | 140 |  |  |  | 20 |  | 0.52702 | 4.5e-07
1 V, 100 KHz |  | 0.99979 | 310 |  |  |  | 30 |  | 1.0002 | 6.5e-07
1 V, 200 KHz |  | 0.99935 | 550 |  |  |  | 70 |  | 1.0145 | 5.7e-07
1 V, 300 KHz |  | 0.99962 | 550 |  |  |  | 70 |  | 0.83365 | 4.7e-07
1 V, 500 KHz |  | 0.99895 | 1400 |  |  |  | 100 |  | 0.50053 | 8.1e-07
1 V, 1 MHz |  | 0.99945 | 3000 |  |  |  | 200 |  | 1.131 | 1.33e-06
10 V, 10 Hz |  | 10.0011 | 600 |  |  |  | 10 |  | 7.2526 | 5.73e-06
10 V, 20 Hz |  | 10.0001 | 180 |  |  |  | 10 |  | 0.99999 | 1.37e-06
10 V, 30 Hz |  | 10 | 180 |  |  |  | 10 |  | 1.1304 | 2.14e-06
10 V, 40 Hz |  | 9.9999 | 77 |  |  |  | 10 |  | 0.70711 | 2.42e-06
10 V, 55 Hz |  | 9.9999 | 77 |  |  |  | 10 |  | 1.965 | 2.23e-06
10 V, 300 Hz |  | 9.9999 | 77 |  |  |  | 10 |  | 1.453 | 1.61e-06
10 V, 1 KHz |  | 10 | 77 |  |  |  | 10 |  | 1.3642 | 1.32e-06
10 V, 10 KHz |  | 9.9998 | 77 |  |  |  | 10 |  | 0.52706 | 3.5e-07
10 V, 20 KHz |  | 10.0001 | 77 |  |  |  | 10 |  | 0.86602 | 4.6e-07
10 V, 30 KHz |  | 10.0001 | 140 |  |  |  | 10 |  | 0.44095 | 3.3e-07
10 V, 50 KHz |  | 10.0003 | 140 |  |  |  | 20 |  | 0.33332 | 3.9e-07
10 V, 100 KHz |  | 9.998 | 270 |  |  |  | 30 |  | 0.8335 | 4.5e-07
10 V, 200 KHz |  | 9.993 | 670 |  |  |  | 70 |  | 0.60134 | 6.8e-07
10 V, 300 KHz |  | 9.996 | 670 |  |  |  | 70 |  | 0.60117 | 4e-07
10 V, 500 KHz |  | 9.9891 | 1700 |  |  |  | 100 |  | 0.86697 | 4.2e-07
10 V, 1 MHz |  | 9.9571 | 3500 |  |  |  | 200 |  | 1.23 | 1.48e-06
19 V, 10 Hz |  | 19.0067 | 552.63 |  |  |  | 10 |  | 7.6836 | 4.5e-06
19 V, 20 Hz |  | 19.0017 | 165.79 |  |  |  | 10 |  | 1.0229 | 1.81e-06
19 V, 30 Hz |  | 19.0008 | 165.79 |  |  |  | 10 |  | 0.64457 | 1.18e-06
19 V, 40 Hz |  | 19.0004 | 73.684 |  |  |  | 10 |  | 1.3099 | 2.15e-06
19 V, 55 Hz |  | 19.0002 | 73.684 |  |  |  | 10 |  | 2.9824 | 2.34e-06
19 V, 300 Hz |  | 19 | 73.684 |  |  |  | 10 |  | 2.3554 | 1.62e-06
19 V, 1 KHz |  | 19.0001 | 73.684 |  |  |  | 10 |  | 1.0563 | 1.23e-06
19 V, 10 KHz |  | 18.9999 | 73.684 |  |  |  | 10 |  | 1.1471 | 4.7e-07
19 V, 20 KHz |  | 19.0004 | 73.684 |  |  |  | 10 |  | 0.5752 | 5.1e-07
19 V, 30 KHz |  | 19.0006 | 130.53 |  |  |  | 10 |  | 0.91158 | 3.9e-07
19 V, 50 KHz |  | 19.0009 | 130.53 |  |  |  | 20 |  | 0.43857 | 2.6e-07
19 V, 100 KHz |  | 18.9967 | 251.05 |  |  |  | 30 |  | 0.87734 | 3.8e-07
19 V, 200 KHz |  | 18.9875 | 589.47 |  |  |  | 70 |  | 1.2568 | 3.5e-07
19 V, 300 KHz |  | 18.9934 | 589.47 |  |  |  | 70 |  | 2.056 | 4e-07
19 V, 500 KHz |  | 18.9804 | 1463.2 |  |  |  | 100 |  | 1.3024 | 4.1e-07
19 V, 1 MHz |  | 18.9234 | 3073.7 |  |  |  | 200 |  | 7.4171 | 1.12e-06
100 V, 10 Hz |  | 100.0108 | 600 |  |  |  | 10 |  | 4.3136 | 1.056e-05
100 V, 20 Hz |  | 100.0011 | 180 |  |  |  | 10 |  | 0.86602 | 1.12e-06
100 V, 30 Hz |  | 99.9996 | 180 |  |  |  | 10 |  | 0.70711 | 2.18e-06
100 V, 40 Hz |  | 99.9987 | 85 |  |  |  | 10 |  | 1.118 | 2.39e-06
100 V, 55 Hz |  | 99.9986 | 85 |  |  |  | 10 |  | 1.1667 | 2.52e-06
100 V, 300 Hz |  | 99.9979 | 85 |  |  |  | 10 |  | 1.9861 | 2.24e-06
100 V, 1 KHz |  | 99.9976 | 85 |  |  |  | 10 |  | 1.1304 | 1.29e-06
100 V, 10 KHz |  | 99.9977 | 85 |  |  |  | 10 |  | 0.66668 | 4.3e-07
100 V, 20 KHz |  | 99.9969 | 85 |  |  |  | 10 |  | 0.50002 | 2e-07
100 V, 30 KHz |  | 99.9977 | 240 |  |  |  | 10 |  | 1.0929 | 7.8e-07
100 V, 50 KHz |  | 99.9963 | 240 |  |  |  | 20 |  | 0.78176 | 7.4e-07
100 V, 100 KHz |  | 99.9905 | 600 |  |  |  | 30 |  | 2.0885 | 7.4e-07
100 V, 200 KHz |  | 99.9852 | 2600 |  |  |  | 50 |  | 3.2408 | 1.71e-06
1000 V, 55 Hz |  | 1000.014 | 79 |  |  |  | 15 |  | 1.3333 | 2.6e-06
1000 V, 300 Hz |  | 1000.023 | 79 |  |  |  | 15 |  | 2.1666 | 2.11e-06
700 V, 1 KHz |  | 699.996 | 80.714 |  |  |  | 50 |  | 1.0102 | 1.35e-06
DCI PERFORMANCE TEST
DCI Test | Expected Value | Measured A | Source Unc. ppm | Lower Limit A | Upper Limit A | Measured ppm | DUT Transfer STB ppm | Test Result % of SPEC | MEAS SDEV ppm
1 µADC | 1e-06 | 9.984e-07 | 10061.0978 | [5 empty] | 0 | 1.229e-05
10 µADC | 1e-05 | 9.9988e-06 | 1045.1254 | [5 empty] | 8.1659 | 1.02e-06
50 µADC | 5e-05 | 5.0001e-05 | 245 | [5 empty] | 1.1547 | 2.5e-07
100 µADC | 0.0001 | 0.0001 | 145 |  |  |  | 7 |  | 0 | 1.5e-07
-100 µADC | -0.0001 | -0.00010001 | 145 |  |  |  | 7 |  | -0.57732 | 1.7e-07
-50 µADC | -5e-05 | -5.0004e-05 | 245 | [5 empty] | -1.1546 | 1.8e-07
-10 µADC | -1e-05 | -1.0002e-05 | -1044.8015 | [5 empty] | 0 | 1.3e-06
0.1 mADC | 0.0001 | 9.9994e-05 | 145 | [5 empty] | 5.0003 | 1.34e-06
0.5 mADC | 0.0005 | 0.0005 | 65 | [5 empty] | 0 | 2.9e-07
1 mADC | 0.001 | 0.001 | 55 |  |  |  | 7 |  | 0 | 1.1e-07
-1 mADC | -0.001 | -0.001 | 55 |  |  |  | 7 |  | -0.57733 | 1.4e-07
-0.5 mADC | -0.0005 | -0.00050002 | 65 | [5 empty] | -1.1547 | 2.3e-07
1 mADC | 0.001 | 0.00099935 | 55 | [5 empty] | 5.0033 | 1.69e-06
5 mADC | 0.005 | 0.0049996 | 65 | [5 empty] | 0 | 2.7e-07
10 mADC | 0.01 | 0.01 | 55 |  |  |  | 7 |  | 0 | 1.7e-07
-10 mADC | -0.01 | -0.010001 | 55 |  |  |  | 7 |  | -0.49997 | 7e-08
-5 mADC | -0.005 | -0.0050007 | 65 | [5 empty] | -0.99986 | 3.2e-07
10 mADC | 0.01 | 0.0099983 | 55 | [5 empty] | 0 | 1.11e-06
50 mADC | 0.05 | 0.049998 | 75 | [5 empty] | 0 | 2.4e-07
100 mADC | 0.1 | 0.099998 | 65 |  |  |  | 7 |  | 0 | 1.1e-07
-100 mADC | -0.1 | -0.1 | 65 |  |  |  | 7 |  | -0.49999 | 1.3e-07
-50 mADC | -0.05 | -0.050001 | 75 | [5 empty] | -1.1547 | 2.5e-07
0.1 ADC | 0.1 | 0.099997 | 65 | [5 empty] | 0 | 1.25e-06
0.5 ADC | 0.5 | 0.49998 | 135 | [5 empty] | 1 | 1.8e-07
-0.5 ADC | -0.5 | -0.50002 | 135 | [5 empty] | -1.1547 | 2.7e-07
-1.0 ADC | -1 | -0.99999 | 105 |  |  |  | 20 |  | 0 | 1.3e-07
1.0 ADC | 1 | 0.99996 | 105 |  |  |  | 20 |  | 0 | 1.2e-07
ACI PERFORMANCE TEST
ACI Test | Expected Value | Measured Aac | Source Unc. ppm | Lower Limit Aac | Upper Limit Aac | Measured ppm | DUT Transfer STB ppm | Test Result % of SPEC | MEAS SDEV ppm
10 µA, 10 Hz | 1e-05 | 1e-05 | 3649.854 | [5 empty] | 78.046 | 1.317e-05
10 µA, 20 Hz | 1e-05 | 1e-05 | 2849.9003 | [5 empty] | 5.7733 | 4.87e-06
10 µA, 30 Hz | 1e-05 | 1e-05 | 2849.886 | [5 empty] | 8.1646 | 5.45e-06
10 µA, 40 Hz | 1e-05 | 1.0001e-05 | 2119.894 | [5 empty] | 9.9995 | 2.41e-06
10 µA 55 Hz | 1e-05 | 1.0001e-05 | 2119.894 | [5 empty] | 4.9998 | 3.76e-06
10 µA, 300 Hz | 1e-05 | 9.9999e-06 | 2120.0212 |  |  | -40 |  |  | 49.9171 | 1.06e-05
10 µA, 1 KHz | 1e-05 | 1.0001e-05 | 2119.8728 | [5 empty] | 4.9997 | 2.98e-05
10 µA, 5 KHz | 1e-05 | 9.9998e-06 | 5500.1375 | [5 empty] | 40.8258 | 9.74e-06
10 µA, 10 KHz | 1e-05 | 9.998e-06 | 11502.2429 | [5 empty] | 14.1449 | 3.39e-06
100 µA, 10 Hz | 0.0001 | 0.00010002 | 950 | 9.9939e-05 | 0.00010006 | -40 | 50 |  | 8.8111 | 7.24e-06
100 µA, 20 Hz | 0.0001 | 0.00010001 | 600 |  |  |  | 50 |  | 0.49995 | 1.66e-06
100 µA, 30 Hz | 0.0001 | 0.00010001 | 600 | 9.9939e-05 | 0.00010006 |  | 50 |  | 0.49996 | 2.19e-06
100 µA, 40 Hz | 0.0001 | 0.00010001 | 320 |  |  |  | 50 |  | 0.95735 | 2.61e-06
100 µA 55 Hz | 0.0001 | 0.00010001 | 320 | 9.9939e-05 | 0.00010006 | -40 | 50 |  | 0.95736 | 2.18e-06
100 µA, 300 Hz | 0.0001 | 0.00010001 | 320 |  |  |  | 50 |  | 4.4345 | 1.7e-06
100 µA, 1 KHz | 0.0001 | 0.0001 | 320 | 9.9939e-05 | 0.00010006 |  | 50 |  | 1.2583 | 3.51e-06
100 µA, 5 KHz | 0.0001 | 0.0001 | 0.1 |  |  |  | 100 |  | 0.49998 | 3.8e-07
100 µA, 10 KHz | 0.0001 | 0.00010002 | 0.25 | 9.9939e-05 | 0.00010006 |  | 300 |  | 0.57723 | 5.5e-07
1mA, 10 Hz | 0.001 | 0.0010002 | 700 |  |  |  | 40 |  | 4.4245 | 1.016e-05
1 mA, 20 Hz | 0.001 | 0.0010001 | 390 | 9.9939e-05 | 0.00010006 |  | 40 |  | 1.1546 | 1.36e-06
1 mA, 30 Hz | 0.001 | 0.0010001 | 390 |  |  |  | 40 |  | 0.57732 | 1.94e-06
1 mA, 40 Hz | 0.001 | 0.001 | 160 | 9.9939e-05 | 0.00010006 |  | 40 |  | 1.1546 | 2.21e-06
1 mA, 55 Hz | 0.001 | 0.0010001 | 160 |  |  |  | 40 |  | 2.0816 | 2.09e-06
1 mA, 300 Hz | 0.001 | 0.001 | 160 | 9.9939e-05 | 0.00010006 |  | 40 |  | 2.4999 | 1.86e-06
1 mA, 1 KHz | 0.001 | 0.001 | 160 |  |  |  | 40 |  | 0.81646 | 1.43e-06
1 mA, 5 KHz | 0.001 | 0.001 | 1000 | 9.9939e-05 | 0.00010006 |  | 70 |  | 0.57733 | 5.7e-07
1 mA, 10 KHz | 0.001 | 0.00099999 | 0.25 |  |  |  | 200 |  | 0.95744 | 4.1e-07
10 mA, 10 Hz | 0.01 | 0.010002 | 700 | 9.9939e-05 | 0.00010006 |  | 40 |  | 4.3485 | 1.047e-05
10 mA, 20 Hz | 0.01 | 0.010001 | 390 |  |  |  | 40 |  | 1.2909 | 1.79e-06
10 mA, 30 Hz | 0.01 | 0.010001 | 390 | 9.9939e-05 | 0.00010006 |  | 40 |  | 0.49997 | 2.63e-06
10 mA, 40 Hz | 0.01 | 0.010001 | 160 |  |  |  | 40 |  | 0.49997 | 2.32e-06
10 mA, 55 Hz | 0.01 | 0.01 | 160 | 9.9939e-05 | 0.00010006 |  | 40 |  | 1.2583 | 2.25e-06
10 mA, 300 Hz | 0.01 | 0.01 | 160 |  |  |  | 40 |  | 1.7077 | 2.03e-06
10 mA, 1 KHz | 0.01 | 0.010001 | 160 | 9.9939e-05 | 0.00010006 |  | 40 |  | 1.2909 | 1.17e-06
10 mA, 5 KHz | 0.01 | 0.010001 | 1000 |  |  |  | 70 |  | 0.49993 | 4e-07
10 mA, 10 KHz | 0.01 | 0.0099999 | 0.25 | 9.9939e-05 | 0.00010006 |  | 200 |  | 1 | 4e-07
100 mA, 10 Hz | 0.1 | 0.10002 | 700 |  |  |  | 40 |  | 10.6864 | 8.33e-06
100 mA, 20 Hz | 0.1 | 0.10001 | 390 | 9.9939e-05 | 0.00010006 |  | 40 |  | 1.7076 | 1.2e-06
100 mA, 30 Hz | 0.1 | 0.10001 | 390 |  |  |  | 40 |  | 0.9999 | 2.86e-06
100 mA, 40 Hz | 0.1 | 0.10001 | 160 | 9.9939e-05 | 0.00010006 |  | 40 |  | 0.5773 | 2.18e-06
100 mA, 55 Hz | 0.1 | 0.10001 | 160 |  |  |  | 40 |  | 0.81643 | 2.3e-06
100 mA, 300 Hz | 0.1 | 0.10001 | 160 | 9.9939e-05 | 0.00010006 |  | 40 |  | 0.95735 | 1.83e-06
100 mA, 1 KHz | 0.1 | 0.10001 | 160 |  |  |  | 40 |  | 0.57729 | 9.1e-07
100 mA, 5 KHz | 0.1 | 0.099997 | 0.1 | 9.9939e-05 | 0.00010006 |  | 70 |  | 0 | 5.9e-07
100 mA, 10 KHz | 0.1 | 0.099991 | 0.25 |  |  |  | 200 |  | 0.50004 | 4.2e-07
1 A, 10 Hz | 1 | 0.99997 | -1 | 9.9939e-05 | 0.00010006 |  | 40 |  | 6.6585 | 1.018e-05
1 A, 20 Hz | 1 | 0.99998 | -1 |  |  |  | 40 |  | 1.5 | 1.78e-06
1 A, 30 Hz | 1 | 0.99997 | -1 | 9.9939e-05 | 0.00010006 |  | 40 |  | 0.57737 | 2.38e-06
1 A, 40 Hz | 1 | 1.0001 | 640 |  |  |  | 40 |  | 1.2582 | 1.47e-06
1 A, 55 Hz | 1 | 1.0001 | 640 | 9.9939e-05 | 0.00010006 |  | 40 |  | 0.81643 | 2.96e-06
1 A, 300 Hz | 1 | 1.0001 | 640 |  |  |  | 40 |  | 1.8928 | 2.03e-06
1 A, 1 KHz | 1 | 1 | 640 | 9.9939e-05 | 0.00010006 |  | 40 |  | 2.0816 | 9.8e-07
1 A, 5 KHz | 1 | 0.99944 | 800 |  |  |  | 70 |  | 0.81695 | 4.6e-07
1 A, 10 KHz | 1 | 1.0001 | 0.82 | 9.9939e-05 | 0.00010006 |  | 200 |  | 0.5773 | 4.7e-07
Test Summary | Initialization | DCV | OHM | ACV | DCI | ACI |  | Total
Time escape minute | 0.19 | 206.93 | 249.15 | 645.88 | 40.4 | 185.6